## Frontpage
Traceability Matrix - ISP13 MR Cardiac Suite
DHF Number: | D000996539
Project Name | AVW15 (ISP13)

## Introduction
Purpose and Scope
This document lists and traces from Design Input to Design Output, Design Verifictaion and Design Validation of AVW15 (ISP13) for MR Cardiac Suite
References
Reference | Identification | Title / additional remarks
Ref 1 | DHF380082 | ISP13 URS
Ref 2 | DHF380211 | ISP13 PRS
Ref 3 | DHF380234 | ISP13 System Design Specification (SDS)
Ref 4 | D000856146 | ISP13 MR Cardiac Suite DRS
Ref 5 | D000879224 | ISP13 MR Cardiac Suite RMM
Ref 6 | D000993081 | ISP13 MR Cardiac Suite DDS
Ref 7 | D000994617 | ISP13 ISP13 MR Cardiac Suite -Viewer Verification Spec
Ref 8 | D000996525 | ISP13 MR Cardiac Suite -Functional LV & RV Verification Spec
Ref 9 | D000996542 | ISP13 MR Cardiac Suite- Functional Long Axis
Ref 10 | D000996541 | ISP13 MR Cardiac Suite -Qflow Verification Spec
Ref 11 | D000996623 | ISP13 MR Cardiac Suite - Spatial Enhancement Verification Spec
Ref 12 | D000996308 | ISP13_Design Verification Report
Ref 13 | DHF380226 | ISP13 Design Validation Plan
Ref 14 | D001104277 | ISP13 Temporal Test Specification
Ref 15 | D001104051 | ISP13 CardiacMapping Test Specification
Terminology & Abbreviations
Terminology & Abbreviations | Description/Definition
URS | User Requirements Specifications
PRS | Product Requirements Specifications
DRS | Detailed Requirements Specifications
SDS | System Design Specification
DDS | Detail Design Specification
VS | Verification Specification
VR | Verification Report
MR | Magnetic Resonance

## Tracebility-MR Viewer
[5 empty] | DESIGN INPUT | [5 empty] | DESIGN VERIFICATION |  |  | DESIGN VALIDATION
SL No | URS Requirement ID | URS Feature (DHF380082,LH) | PRS Requirement ID | PRS Feature (DHF380211,Rev LJ ) | SDS ID (DHF380234 ,Rev LE) | DRS Requirement ID | DRS Feature (D000856146,Rev I) | RMM ID (D000879224 , Rev F) | DDS Id (D000993081,Rev D) | Verification Test Case (ID / Brief Description) (D000994617,Rev G) | Verification Test Step ID (D000994617,Rev G) | Verification Test Result (PASS/Fail) [Verification_Report: D000996308, Rev E] Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | Defect ID (for Failed steps from Verification) | Validation Test ID (Validation Plan DHF380226, Rev D, Validation Report D000926820 , Rev C) | Validation Test Result (Pass/Fail) [Validation Report D000926820, Rev C] | Defect ID (for Failed steps from Validation) | Remarks
1 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353723 | Common: incorrect patient name or patient ID in Patient Bar [4.1.1][4.1.2] | Cardiac Suite_Viewer_Common RMM 1353723: 4.1.1 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_01 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | Pass | None
2 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353723 | Common: incorrect patient name or patient ID in Patient Bar [4.1.1][4.1.2] | Cardiac Suite_Viewer_Common RMM 1353723: 4.1.2 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_02 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
3 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353745 | Common: Wrong additional patient information in the Rich Patient Demographic: Patient Name, Patient ID, Date of Birth, Referring Physician, Referring Physician Contact details. [4.1.3] | Cardiac Suite_Viewer_Common RMM 1353745: 4.1.3 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_03 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
4 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353745 | Common: Wrong additional patient information in the Rich Patient Demographic: Patient Name, Patient ID, Date of Birth, Referring Physician, Referring Physician Contact details. [4.1.3] | Cardiac Suite_Viewer_Common RMM 1353745: 4.1.3 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_04 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
5 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353745 | Common: Wrong additional patient information in the Rich Patient Demographic: Patient Name, Patient ID, Date of Birth, Referring Physician, Referring Physician Contact details. [4.1.3] | Cardiac Suite_Viewer_Common RMM 1353745: 4.1.3 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_05 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
6 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353785 | Common: Date format [4.2.1] | Cardiac Suite_Viewer_Common RMM 1353785: 4.2.1 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_06 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
7 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353829 | Common: Wrong exam information in Timeline: on the exam thumbnail incorrect information is displayed.[4.2.2] | Cardiac Suite_Viewer_Common RMM 1353829: 4.2.2 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_07 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
8 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353790 | Common: Wrong exam information in Timeline: on the series thumbnail incorrect series number, series description, modality or number of images information is displayed.[4.2.3] | Cardiac Suite_Viewer_Common RMM 1353790: 4.2.3 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_08 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
9 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353789 | Common: Wrong exam information in Timeline: incorrect reference between selected exam thumbnail and presented exam content (shows e.g. series, key images, findings of a wrong exam).[4.2.4] | Cardiac Suite_Viewer_Common RMM 1353789: 4.2.4 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_09 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
10 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353789 | Common: Wrong exam information in Timeline: incorrect reference between selected exam thumbnail and presented exam content (shows e.g. series, key images, findings of a wrong exam).[4.2.4] | Cardiac Suite_Viewer_Common RMM 1353789: 4.2.4 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
11 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353791 | Common: Wrong exam information: when there is a filtering applied on the exams, wrong or irrelevant exams are still presented. [4.2.5] | Cardiac Suite_Viewer_Common RMM 1353791: 4.2.5 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
12 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353808 | Common: Wrong exam information: when there is a filtering applied on the exams, relevant exams are being filtered. [4.2.6] | Cardiac Suite_Viewer_Common RMM 1353808: 4.2.6 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
13 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355337 | Common: Incorrect Study Date is displayed on the images.[4.2.23] | Cardiac Suite_Viewer_Common RMM 1355337: 4.2.23 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
14 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353780 | Common: The number of exams displayed in the Timebar does not match the number of exams available in the Docked Timeline. [4.1.11] | Cardiac Suite_Viewer_Common RMM 1353780: 4.1.11 | 4.3.8.Cardiac (MR) Viewer | 1450151 | Viewer1_Haz_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
15 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1353819 | Common: The user is associating a finding with the wrong exam [4.2.16] | Cardiac Suite_Viewer_Common RMM 1353819: 4.2.16 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_01 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
16 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353819 | Common: The user is associating a finding with the wrong exam [4.2.16] | Cardiac Suite_Viewer_Common RMM 1353819: 4.2.16 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_02 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
17 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353819 | Common: The user is associating a finding with the wrong exam [4.2.16] | Cardiac Suite_Viewer_Common RMM 1353819: 4.2.16 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_03 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
18 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353819 | Common: The user is associating a finding with the wrong exam [4.2.16] | Cardiac Suite_Viewer_Common RMM 1353819: 4.2.16 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_04 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
19 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1353826 | Common: The series was split. The user views one of the series the original series was split into. The user thinks that information is missing. User orders another scan. [4.2.20] | Cardiac Suite_Viewer_Common RMM 1353826: 4.2.20 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_05 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
20 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355340 | Common: Reversing of L/R, A/P, H/F. 2. Incorrect volume orientation (axial/coronal/sagittal, orientation box).[4.3.1] | Cardiac Suite_Viewer_Common RMM 1355340: 4.3.1 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_06 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
21 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355340 | Common: Reversing of L/R, A/P, H/F. 2. Incorrect volume orientation (axial/coronal/sagittal, orientation box).[4.3.1] | Cardiac Suite_Viewer_Common RMM 1355340: 4.3.1 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_07 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
22 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355342 | Common: Incorrect display of Pixel data due to 2D image manipulations (e.g. zoom, pan, windowing, flip).[4.3.2] | Cardiac Suite_Viewer_Common RMM 1355342: 4.3.2 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_08 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
23 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355344 | Common: Wrong or missing contrast indication on displayed image.[4.3.4] | Cardiac Suite_Viewer_Common RMM 1355344: 4.3.4 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_09 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
24 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355344 | Common: Wrong or missing contrast indication on displayed image.[4.3.4] | Cardiac Suite_Viewer_Common RMM 1355344: 4.3.4 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
25 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355352 | Common: The user is scrolling images in a series using mouse click and drag. The user is scrolling in a speed that causes the system to skip slices. The user is unaware that Not all images have been shown.[4.3.12] | Cardiac Suite_Viewer_Common RMM 1355352: 4.3.12 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
26 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355353 | Common: User loaded 2 series with different field of view. The “zoom link” is causing a part of the anatomy to be out of view The user think he is looking at the image in the original zoom.[4.3.14] | Cardiac Suite_Viewer_Common RMM 1355353: 4.3.14 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
27 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355354 | Common: Showing wrong slice number under "phase" annotation in addition to the correct slice number under the "slice" annotation [4.3.15] | Cardiac Suite_Viewer_Common RMM 1355354: 4.3.15 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
28 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238807 | Link and Sync: Relate (Localizer) | NA | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
29 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238807 | Link and Sync: Relate (Localizer) | NA | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
30 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357574 | Common: Inaccurate automatic processing results (bone removal, center-line extraction, contour extraction, caliper annotation) which is being implicitly applied without user's interference[4.4.1 ] | Cardiac Suite_Viewer_Common RMM 1357574: 4.4.1 | 4.3.8.Cardiac (MR) Viewer | 1450153 | Viewer2_Haz_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
31 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355629 | Common: Geometrical measurement value is displayed with wrong units or without units[4.4.5.1 ] | 1355629: 4.4.5.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_01 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
32 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355355 | Common: Line measurements show wrong value with more than allowed error[4.4.2.1] | 1355355: 4.4.2.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_01 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
33 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355355 | Common: Line measurements show wrong value with more than allowed error[4.4.2.1] | 1355355: 4.4.2.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_02 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
34 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355629 | Common: Geometrical measurement value is displayed with wrong units or without units[4.4.5.1 ] | 1355629: 4.4.5.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_02 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
35 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355355 | Common: Line measurements show wrong value with more than allowed error[4.4.2.1] | 1355355: 4.4.2.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_03 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
36 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355629 | Common: Geometrical measurement value is displayed with wrong units or without units[4.4.5.1 ] | 1355629: 4.4.5.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_03 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
37 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355357 | Common: Angle measurements show wrong value with more than allowed error[4.4.2.2] | 1355357: 4.4.2.2 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_04 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
38 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355629 | Common: Geometrical measurement value is displayed with wrong units or without units[4.4.5.1 ] | 1355629: 4.4.5.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_04 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
39 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355357 | Common: Angle measurements show wrong value with more than allowed error[4.4.2.2] | 1355357: 4.4.2.2 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_05 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
40 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355629 | Common: Geometrical measurement value is displayed with wrong units or without units[4.4.5.1 ] | 1355629: 4.4.5.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_05 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
41 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355357 | Common: Angle measurements show wrong value with more than allowed error[4.4.2.2] | 1355357: 4.4.2.2 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_06 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
42 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355629 | Common: Geometrical measurement value is displayed with wrong units or without units[4.4.5.1 ] | 1355629: 4.4.5.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_06 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
43 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355595 | Common: 2D ROI measurements show wrong value with more than allowed error[4.4.2.3] | 1355595: 4.4.2.3 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_07 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
44 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355629 | Common: Geometrical measurement value is displayed with wrong units or without units[4.4.5.1 ] | 1355629: 4.4.5.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_07 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
45 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355595 | Common: 2D ROI measurements show wrong value with more than allowed error[4.4.2.3] | 1355595 : 4.4.2.3 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_08 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
46 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355629 | Common: Geometrical measurement value is displayed with wrong units or without units[4.4.5.1 ] | 1355629 : 4.4.5.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_08 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
47 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355601 | Common: 3D volume measurements show wrong value with more than allowed error[4.4.2.4 ] | 1355601 : 4.4.2.4 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_09 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
48 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355629 | Common: Geometrical measurement value is displayed with wrong units or without units[4.4.5.1 ] | 1355629 : 4.4.5.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_09 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
49 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355641 | Common: Line measurements units[4.4.5.2 ] | 1355641 : 4.4.5.2 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
50 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355999 | Common: Angle measurements units[4.4.5.3 ] | 1355999 : 4.4.5.3 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
51 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1355641 | Common: Line measurements units[4.4.5.2 ] | 1355641 : 4.4.5.2 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
52 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1356000 | Common: ROI (Area) measurements units[4.4.5.4 ] | 1356000 : 4.4.5.4 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
53 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1356000 | Common: ROI (Area) measurements units[4.4.5.4 ] | 1356000 : 4.4.5.4 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
54 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1356001 | Common: VOI (Volume ) measurements units [4.4.5.5 ] | 1356001 : 4.4.5.5 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
55 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1356009 | Common: Geometrical measurement value is changed when saving a finding[4.4.6 ] | 1356009 : 4.4.6 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
56 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1356011 | Common: Geometrical measurement value or position is changed after activating a saved finding[4.4.7 ] | 1356011 : 4.4.7 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
57 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 1356013 : 4.4.8 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
58 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 1356013 : 4.4.8 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
59 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357584 | Common: Pixel : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness)[4.4.9.5] | 1357584: 4.4.9.5 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
60 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357583 | Common: Angle: Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness)[4.4.9.4 ] | 1357583: 4.4.9.4 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
61 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357582 | Common: Line : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness)[4.4.9.3 ] | 1357582 : 4.4.9.3 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
62 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1356014 | Common: Geometrical measurements' values changed with more than allowed error due to image manipulations ( pan, rotate, mirror, invert, sharpness)[4.4.9.1 ] | 1356014 : 4.4.9.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
63 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357584 | Common: Pixel : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness)[4.4.9.5] | 1357584: 4.4.9.5 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_19 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
64 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357582 | Common: Line : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness)[4.4.9.3 ] | 1357582 : 4.4.9.3 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_19 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
65 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357583 | Common: Angle: Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness)[4.4.9.4 ] | 1357583: 4.4.9.4 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_19 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
66 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1356014 | Common: Geometrical measurements' values changed with more than allowed error due to image manipulations ( pan, rotate, mirror, invert, sharpness)[4.4.9.1 ] | 1356014 : 4.4.9.1 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
67 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357582 | Common: Line : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness)[4.4.9.3 ] | 1357582 : 4.4.9.3 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
68 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357583 | Common: Angle: Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness)[4.4.9.4 ] | 1357583: 4.4.9.4 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
69 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357584 | Common: Pixel : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness)[4.4.9.5] | 1357584: 4.4.9.5 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
70 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357586 | Common: VOI: Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness) [4.4.9.6] | 1357586 : 4.4.9.6 | 4.3.8.Cardiac (MR) Viewer | 1450155 | Viewer3_Measurent_20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
71 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357587 | Common: Sharpness : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness) [4.4.9.7] | 1357587 : 4.4.9.7 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_01 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
72 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357587 | Common: Sharpness : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness) [4.4.9.7] | 1357587 : 4.4.9.7 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_02 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
73 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357588 | Common: Thickenss: Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness) [4.4.10] | 1357588 : 4.4.10 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_03 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
74 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357589 | Common: Incorrect displayed Pixel Values on the images [ 4.4.11] | 1357589 : 4.4.11 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_04 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
75 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357590 | Common: Pixel Value is displayed with wrong units [ 4.4.12] | 1357590 : 4.4.12 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_04 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
76 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357605 | Common: PT SUV values displayed without indication on the used calculation method [ 4.4.14 ] | 1357605 : 4.4.14 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_05 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
77 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357605 | Common: PT SUV values displayed without indication on the used calculation method [ 4.4.14 ] | 1357605 : 4.4.14 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_06 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
78 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357606 | Common: Incorrect or incorrect presentation of scale, grid on images.[ 4.4.15] | 1357606 : 4.4.15 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_07 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
79 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357610 | Common: Textual annotation's position is changed due to image manipulations (e.g. zoom, pan, sweivel, rotate, flip)[ 4.4.20] | 1357610: 4.4.20 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_08 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
80 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357607 | Common: The user performed a measurement unsuccessfully, the user then wants to correct the measurement It is too hard for the user or the user don’t understand how to edit an incorrect measurement.[ 4.4.16] | 1357607 : 4.4.16 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_08 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
81 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357613 | Common: inding preview image is wrong or inaccurate[ 4.4.26] | 1357613 : 4.4.26 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
82 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357614 | Common: Loading by default with non-original thickness in Slab mode and the user is unaware that the thickness is not at scanned minimum[ 4.4.28] | 1357614 : 4.4.28 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
83 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357615 | Common: The user segmented a lesion The user didn’t verify the segmentation on all of the slices. The segmentation is not accurate, leading to misleading results[ 4.4.30] | Cardiac Suite_Viewer_Common RMM 1357615 : 4.4.30 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
84 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357617 | Common: The user segmented a lesion. The user edits a segmentation by mistake leading to wrong results.s[ 4.4.31] | 1357617 : 4.4.31 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
85 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357618 | Common: The user segmented a lesion Long and short axis doesn’t appear on the segmentation The user can’t verify how the Long and short axis where measured[ 4.4.32] | 1357618 : 4.4.32 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
86 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357620 | Common: Wrong interpretation of images when lossy compression is used 2. Wrong interpretation of images when lossy compression with resolution reduction is used[ 4.5.1] | 1357620 : 4.5.1 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
87 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357621 | Common: wrong interpretation of images when the view mode is set to "Slab"[ 4.5.2] | 1357621 : 4.5.2 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
88 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1361982 | Common: mixed patients [4.6.1] | 1361982 : 4.6.1 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
89 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357862 | Common: Identify between multiple studies.[ 4.6.1.1] | 1357862 : 4.6.1.1 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
90 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1362074 | Common: new images are saved to an exam of the wrong patient.[4.6.2] | 1362074 : 4.6.2 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
91 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357871 | Common: new images are saved to a wrong exam of the same patient[ 4.7.1] | 1357871 : 4.7.1 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_20 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
92 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357885 | Common: the letter and color indication of the exam (tagging icon) on the viewport is incorrect[ 4.7.2] | 1357885 : 4.7.2 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_21 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
93 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357885 | Common: the letter and color indication of the exam (tagging icon) on the viewport is incorrect[ 4.7.2] | 1357885 : 4.7.2 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_22 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
94 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357885 | Common: the letter and color indication of the exam (tagging icon) on the viewport is incorrect[ 4.7.2] | 1357885 : 4.7.2 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_23 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
95 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357885 | Common: the letter and color indication of the exam (tagging icon) on the viewport is incorrect[ 4.7.2] | 1357885 : 4.7.2 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_24 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
96 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate..[ 4.8.1, 4.8.2, 4.8.3] | 1357886 : 4.8.3 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_25 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
97 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate..[ 4.8.1, 4.8.2, 4.8.3] | 1357886 : 4.8.1 | 4.3.8.Cardiac (MR) Viewer | 1450157 | Viewer4_Haz_25 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
98 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357891 | Common: The saved finding doesn't appear where user expects it [ 4.8.7] | Cardiac Suite_Viewer_Common RMM 1357891 : 4.8.7 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_01 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
99 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357894 | Common: The user selects “All series combined” from the series widget in the header of the arrangement. The user doesn’t know which series appears in the arrangement.[ 4.8.10] | Cardiac Suite_Viewer_Common RMM 1357894 : 4.8.10 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_02 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
100 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1359246 | Common: Viewing or Analysis based on incorrect or wrong merging of split series [4.10] | Cardiac Suite_Viewer_Common RMM 1359246 : 4.10 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_03 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
101 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357594 | Common: incorrect geometrical measurements on images due to loss or change in calibration factor: DX, CR, XA, RF [ 4.4.13.1] | Cardiac Suite_Viewer_Common RMM 1357594 : 4.4.13.1 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_05 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
102 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357594 | Common: incorrect geometrical measurements on images due to loss or change in calibration factor: DX, CR, XA, RF [ 4.4.13.1] | Cardiac Suite_Viewer_Common RMM 1357594 : 4.4.13.1 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_06 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
103 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357594 | Common: incorrect geometrical measurements on images due to loss or change in calibration factor: DX, CR, XA, RF [ 4.4.13.1] | Cardiac Suite_Viewer_Common RMM 1357594 : 4.4.13.1 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_07 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
104 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357594 | Common: incorrect geometrical measurements on images due to loss or change in calibration factor: DX, CR, XA, RF [ 4.4.13.1] | Cardiac Suite_Viewer_Common RMM 1357594 : 4.4.13.1 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_08 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
105 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357594 | Common: incorrect geometrical measurements on images due to loss or change in calibration factor: DX, CR, XA, RF [ 4.4.13.1] | Cardiac Suite_Viewer_Common RMM 1357594 : 4.4.13.2 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_09 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
106 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357594 | Common: incorrect geometrical measurements on images due to loss or change in calibration factor: DX, CR, XA, RF [ 4.4.13.1] | Cardiac Suite_Viewer_Common RMM 1357594 : 4.4.13.2 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
107 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357594 | Common: incorrect geometrical measurements on images due to loss or change in calibration factor: DX, CR, XA, RF [ 4.4.13.1] | Cardiac Suite_Viewer_Common RMM 1357594 : 4.4.13.2 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Presentation state functionality is not supported in MR Viewer
108 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357594 | Common: incorrect geometrical measurements on images due to loss or change in calibration factor: DX, CR, XA, RF [ 4.4.13.1] | Cardiac Suite_Viewer_Common RMM 1357594 : 4.4.13.2 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Presentation state functionality is not supported in MR Viewer
109 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357594 | Common: incorrect geometrical measurements on images due to loss or change in calibration factor: DX, CR, XA, RF [ 4.4.13.1] | Cardiac Suite_Viewer_Common RMM 1357594 : 4.4.13.2 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Presentation state functionality is not supported in MR Viewer
110 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357888 | Common: some of the images of the series are not presented in the viewer in Slab viewing mode (happens when there are different types of images in the series, or irregular gaps)[ 4.8.4] | Cardiac Suite_Viewer_Common RMM 1357888 : 4.8.4 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
111 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357889 | Common: when soft-deleting images (findings) the wrong images are marked for soft delete.[ 4.8.5] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
112 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357889 | Common: when soft-deleting images (findings) the wrong images are marked for soft delete.[ 4.8.5] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
113 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357889 | Common: when soft-deleting images (findings) the wrong images are marked for soft delete.[ 4.8.5] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
114 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357889 | Common: when soft-deleting images (findings) the wrong images are marked for soft delete.[ 4.8.5] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1450158 | Viewer5_Haz_18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
115 | 1179490 | Suite Workflow | 1182563 1228262 | Finding dashboard (managment) General Behavior | Section 2.2.1 and Section 2.2.2 | 1357891 | Common: The saved finding doesn't appear where user expects it [ 4.8.7] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366755 | MRC_RMF_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
116 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1300189 | Send results to finding dashboard - usability safety mitigations | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366755 | MRC_RMF_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
117 | 1179490 | Suite Workflow | 1182563 | Viewer | Section 2.2.1 and Section 2.2.2 | 1300189 | Send results to finding dashboard - usability safety mitigations | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366755 | MRC_RMF_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
118 | 1179490 | Suite Workflow | 1182563 1228262 | Viewer | Section 2.2.1 and Section 2.2.2 | 1357891 | Common: The saved finding doesn't appear where user expects it [ 4.8.7] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366755 | MRC_RMF_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
119 | 1179490 | Suite Workflow | 1182563 1228262 | Finding dashboard (managment) General Behavior | Section 2.2.1 and Section 2.2.2 | 1357613 | Common: Finding preview image is wrong or inaccurate[ 4.4.26] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366755 | MRC_RMF_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
120 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355342 | Common: Incorrect display of Pixel data due to 2D image manipulations (e.g. zoom, pan, windowing, flip).[4.3.2] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367569 | Common_RMM_01 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
121 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355344 | Common: Wrong or missing contrast indication on displayed image.[4.3.4] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367569 | Common_RMM_02 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
122 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355353 | Common: User loaded 2 series with different field of view. The “zoom link” is causing a part of the anatomy to be out of view The user think he is looking at the image in the original zoom.[4.3.14] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367569 | Common_RMM_03 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
123 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355354 | Common: Showing wrong slice number under "phase" annotation in addition to the correct slice number under the "slice" annotation [4.3.15] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367569 | Common_RMM_04 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
124 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355352 | Common: The user is scrolling images in a series using mouse click and drag. The user is scrolling in a speed that causes the system to skip slices. The user is unaware that Not all images have been shown.[4.3.12] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367569 | Common_RMM_04 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
125 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355352 | Common: The user is scrolling images in a series using mouse click and drag. The user is scrolling in a speed that causes the system to skip slices. The user is unaware that Not all images have been shown.[4.3.12] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367569 | Common_RMM_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
126 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357871 | Common: new images are saved to a wrong exam of the same patient[ 4.7.1] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367400 | Generic_AllApps_RMF_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
127 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate..[ 4.8.1, 4.8.2, 4.8.3] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367400 | Generic_AllApps_RMF_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
128 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate..[ 4.8.1, 4.8.2, 4.8.3] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367400 | Generic_AllApps_RMF_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
129 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate..[ 4.8.1, 4.8.2, 4.8.3] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367400 | Generic_AllApps_RMF_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
130 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate..[ 4.8.1, 4.8.2, 4.8.3] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1367400 | Generic_AllApps_RMF_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
131 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357606 | Common: Incorrect or incorrect presentation of scale, grid on images.[ 4.4.15] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366785 | MRC_CommonRMF_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
132 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357610 | Common: Textual annotation's position is changed due to image manipulations (e.g. zoom, pan, sweivel, rotate, flip)[ 4.4.20] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366785 | MRC_CommonRMF_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
133 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357620 | Common: Wrong interpretation of images when lossy compression is used 2. Wrong interpretation of images when lossy compression with resolution reduction is used[ 4.5.1] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366785 | MRC_CommonRMF_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
134 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357587 | Common: Sharpness : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness) [4.4.9.7] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366785 | MRC_CommonRMF_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
135 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357607 | Common: Sharpness : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness) [4.4.9.7] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366785 | MRC_CommonRMF_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
136 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357610 | Common: Textual annotation's position is changed due to image manipulations (e.g. zoom, pan, sweivel, rotate, flip)[ 4.4.20] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366785 | MRC_CommonRMF_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
137 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357606 | Common: Incorrect or incorrect presentation of scale, grid on images.[ 4.4.15] | Cardiac Suite_Viewer_Common RMM 1357889 : 4.8.5 | 4.3.8.Cardiac (MR) Viewer | 1366785 | MRC_CommonRMF_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
138 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235587 | Viewing: Activate over Hover | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
139 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235590 | Viewing: Fit to monitor | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
140 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235590 | Viewing: Fit to monitor | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
141 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235587 | Viewing: Activate over Hover | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
142 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235587 | Viewing: Activate over Hover | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
143 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235595 | Viewing: Export Image as JPEG | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_6 | Fail Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | 1531798 | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
144 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235595 | Viewing: Export Image as JPEG | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
145 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235590 | Viewing: Fit to monitor | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
146 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236005 | Viewing: Basic Viewing Capabilities | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
147 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235590 | Viewing: Fit to monitor | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
148 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236005 | Viewing: Basic Viewing Capabilities | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
149 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235590 | Viewing: Fit to monitor | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
150 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236005 | Viewing: Basic Viewing Capabilities | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
151 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235590 | Viewing: Fit to monitor | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
152 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236776 | Viewing: "View" menu | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
153 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235590 | Viewing: Fit to monitor | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
154 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235590 | Viewing: Fit to monitor | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
155 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236005 | Viewing: Basic Viewing Capabilities | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
156 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236779 | Viewing: Context menu functionality | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_capabilites_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
157 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236779 | Viewing: Context menu functionality | NA | 4.3.8.Cardiac (MR) Viewer | 1418143 | Context_Menu_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
158 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236779 | Viewing: Context menu functionality | NA | 4.3.8.Cardiac (MR) Viewer | 1418143 | Context_Menu_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
159 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236779 | Viewing: Context menu functionality | NA | 4.3.8.Cardiac (MR) Viewer | 1418143 | Context_Menu_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
160 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236779 | Viewing: Context menu functionality | NA | 4.3.8.Cardiac (MR) Viewer | 1418143 | Context_Menu_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
161 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236779 | Viewing: Context menu functionality | NA | 4.3.8.Cardiac (MR) Viewer | 1418143 | Context_Menu_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
162 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236779 | Viewing: Context menu functionality | NA | 4.3.8.Cardiac (MR) Viewer | 1418143 | Context_Menu_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
163 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236779 | Viewing: Context menu functionality | NA | 4.3.8.Cardiac (MR) Viewer | 1418143 | Context_Menu_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
164 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236779 | Viewing: Context menu functionality | NA | 4.3.8.Cardiac (MR) Viewer | 1418143 | Context_Menu_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
165 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238741 | Viewing: Layout Management | NA | 4.3.8.Cardiac (MR) Viewer | 1418145 | Layout_Max_Quick_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
166 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238741 | Viewing: Layout Management | NA | 4.3.8.Cardiac (MR) Viewer | 1418145 | Layout_Max_Quick_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
167 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238741 | Viewing: Layout Management | NA | 4.3.8.Cardiac (MR) Viewer | 1418145 | Layout_Max_Quick_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
168 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238741 | Viewing: Layout Management | NA | 4.3.8.Cardiac (MR) Viewer | 1418145 | Layout_Max_Quick_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
169 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236780 | Viewing: Maximize\ Minimize | NA | 4.3.8.Cardiac (MR) Viewer | 1418145 | Layout_Max_Quick_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
170 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238746 | Viewing: Quick Series Selection from arrangement header | NA | 4.3.8.Cardiac (MR) Viewer | 1418145 | Layout_Max_Quick_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
171 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238746 | Viewing: Quick Series Selection from arrangement header | NA | 4.3.8.Cardiac (MR) Viewer | 1418145 | Layout_Max_Quick_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
172 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238746 | Viewing: Quick Series Selection from arrangement header | NA | 4.3.8.Cardiac (MR) Viewer | 1418145 | Layout_Max_Quick_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
173 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238746 | Viewing: Quick Series Selection from arrangement header | NA | 4.3.8.Cardiac (MR) Viewer | 1418145 | Layout_Max_Quick_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
174 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238746 | Viewing: Quick Series Selection from arrangement header | NA | 4.3.8.Cardiac (MR) Viewer | 1418145 | Layout_Max_Quick_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
175 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238761 | Viewing: Swap Arrangements | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
176 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238761 | Viewing: Swap Arrangements | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
177 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238761 | Viewing: Swap Arrangements | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
178 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238761 | Viewing: Swap Arrangements | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
179 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238784 | Viewing: Image Annotations | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
180 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238784 | Viewing: Image Annotations | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
181 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238784 | Viewing: Image Annotations | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
182 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238784 | Viewing: Image Annotations | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
183 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238784 | Viewing: Image Annotations | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
184 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238784 | Viewing: Image Annotations | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
185 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238784 | Viewing: Image Annotations | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
186 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238784 | Viewing: Image Annotations | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
187 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238784 | Viewing: Image Annotations | NA | 4.3.8.Cardiac (MR) Viewer | 1418146 | Swap_annotation_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
188 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
189 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
190 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
191 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
192 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
193 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
194 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
195 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_8 | NA | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | In 2D mode image cannot be swiveled,
196 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
197 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
198 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
199 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
200 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
201 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_14 | NA | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | In 2D mode image cannot be swiveled
202 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
203 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
204 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
205 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238790 | Viewing: 3D Mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418159 | Viewing_3D_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
206 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238790 | Viewing: 3D Mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418159 | Viewing_3D_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
207 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238790 | Viewing: 3D Mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418159 | Viewing_3D_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
208 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238790 | Viewing: 3D Mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418159 | Viewing_3D_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
209 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238790 | Viewing: 3D Mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418159 | Viewing_3D_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
210 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238790 | Viewing: 3D Mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418159 | Viewing_3D_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
211 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238790 | Viewing: 3D Mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418159 | Viewing_3D_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
212 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238790 | Viewing: 3D Mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418159 | Viewing_3D_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
213 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238796 | Viewing: Changing viewing mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418160 | Changing_Viewing_Mode_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
214 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238796 | Viewing: Changing viewing mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418160 | Changing_Viewing_Mode_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
215 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238796 | Viewing: Changing viewing mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418160 | Changing_Viewing_Mode_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
216 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238796 | Viewing: Changing viewing mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418160 | Changing_Viewing_Mode_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
217 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238796 | Viewing: Changing viewing mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418160 | Changing_Viewing_Mode_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
218 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238796 | Viewing: Changing viewing mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418160 | Changing_Viewing_Mode_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
219 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238796 | Viewing: Changing viewing mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418160 | Changing_Viewing_Mode_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
220 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
221 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
222 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
223 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
224 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
225 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
226 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
227 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
228 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
229 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
230 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
231 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238804 | Measurements: Show/Hide measurements | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
232 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
233 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238804 | Measurements: Show/Hide measurements | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
234 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
235 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
236 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
237 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
238 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
239 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
240 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_19 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
241 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.8.Cardiac (MR) Viewer | 1418161 | Measurement_20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
242 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
243 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
244 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
245 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
246 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
247 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
248 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
249 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
250 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
251 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
252 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
253 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
254 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
255 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
256 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
257 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
258 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238801 | Measurements: 3D segmentation on CT & MR | NA | 4.3.8.Cardiac (MR) Viewer | 1418162 | 3D_Seg_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
259 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
260 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_2 | PASS | 1501151 | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
261 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
262 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
263 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_5 | PASS | 1501151 | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
264 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
265 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
266 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
267 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239036 | Cine General Flow | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
268 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
269 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239036 | Cine General Flow | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
270 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
271 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
272 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239036 | Cine General Flow | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
273 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239036 | Cine General Flow | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
274 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239036 | Cine General Flow | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
275 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239036 | Cine General Flow | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
276 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239036 | Cine General Flow | NA | 4.3.8.Cardiac (MR) Viewer | 1418163 | Cine_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
277 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240001 | Filtering Panel | NA | 4.3.8.Cardiac (MR) Viewer | 1418176 | Filter_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
278 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240001 | Filtering Panel | NA | 4.3.8.Cardiac (MR) Viewer | 1418176 | Filter_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
279 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240001 | Filtering Panel | NA | 4.3.8.Cardiac (MR) Viewer | 1418176 | Filter_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
280 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240002 | Color Coding | NA | 4.3.8.Cardiac (MR) Viewer | 1418177 | Color_coding_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
281 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240002 | Color Coding | NA | 4.3.8.Cardiac (MR) Viewer | 1418177 | Color_coding_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
282 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240002 | Color Coding | NA | 4.3.8.Cardiac (MR) Viewer | 1418177 | Color_coding_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
283 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240018 | Time bar UI | NA | 4.3.8.Cardiac (MR) Viewer | 1418181 | Timebar_UI_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
284 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240018 | Time bar UI | NA | 4.3.8.Cardiac (MR) Viewer | 1418181 | Timebar_UI_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
285 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240018 | Time bar UI | NA | 4.3.8.Cardiac (MR) Viewer | 1418181 | Timebar_UI_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
286 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240018 | Time bar UI | NA | 4.3.8.Cardiac (MR) Viewer | 1418181 | Timebar_UI_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
287 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240018 | Time bar UI | NA | 4.3.8.Cardiac (MR) Viewer | 1418181 | Timebar_UI_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
288 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239491 | Full Exam Population | NA | 4.3.8.Cardiac (MR) Viewer | 1418181 | Timebar_UI_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
289 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240009 | Dragging exam from timeline to viewer | NA | 4.3.8.Cardiac (MR) Viewer | 1418183 | Exam_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
290 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240009 | Dragging exam from timeline to viewer | NA | 4.3.8.Cardiac (MR) Viewer | 1418183 | Exam_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
291 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240024 | Exam representation | NA | 4.3.8.Cardiac (MR) Viewer | 1418183 | Exam_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
292 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240024 | Exam representation | NA | 4.3.8.Cardiac (MR) Viewer | 1418183 | Exam_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
293 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240010 | Exam Selection | NA | 4.3.8.Cardiac (MR) Viewer | 1418183 | Exam_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
294 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240010 | Exam Selection | NA | 4.3.8.Cardiac (MR) Viewer | 1418183 | Exam_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
295 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240027 | Series representation | NA | 4.3.8.Cardiac (MR) Viewer | 1418183 | Exam_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
296 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239484 | All Series combined | NA | 4.3.8.Cardiac (MR) Viewer | 1418183 | Exam_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
297 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239483 | View next / previous group of series | NA | 4.3.8.Cardiac (MR) Viewer | 1418183 | Exam_9 | Fail | 1500299 | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Defect no paln to fix
298 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240028 | Time Line: Dialog basic UI | NA | 4.3.8.Cardiac (MR) Viewer | 1418185 | Dialog_UI_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
299 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240028 | Time Line: Dialog basic UI | NA | 4.3.8.Cardiac (MR) Viewer | 1418185 | Dialog_UI_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
300 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240028 | Time Line: Dialog basic UI | NA | 4.3.8.Cardiac (MR) Viewer | 1418185 | Dialog_UI_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
301 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240028 | Time Line: Dialog basic UI | NA | 4.3.8.Cardiac (MR) Viewer | 1418185 | Dialog_UI_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
302 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235587 | Viewing: Activate over Hover | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_Viewing_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
303 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240129 | DXR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
304 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240151 | MR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
305 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240129 | DXR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
306 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235587 | Viewing: Activate over Hover | NA | 4.3.8.Cardiac (MR) Viewer | 1418142 | Basic_Viewing_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
307 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235587 | Viewing: Activate over Hover | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
308 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240129 | Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
309 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240151 | MR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
310 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235587 | Viewing: Activate over Hover | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
311 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240129 | Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
312 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240151 | MR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
313 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240151 | MR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
314 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240151 | MR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
315 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240151 | MR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_6 | PASS | 1531798 | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
316 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240151 | MR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
317 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235590 | Viewing: Fit to monitor | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
318 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235590 | Viewing: Fit to monitor | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
319 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240174 | MR: Multi-Dimensional viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
320 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240174 | MR: Multi-Dimensional viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
321 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240115 | Viewing: Multi-Vendor: Basic Viewing Capabilities | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
322 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240129 | DXR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
323 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240129 | Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
324 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1240151 | MR Basic Viewing | NA | 4.3.8.Cardiac (MR) Viewer | 1418190 | Basic_Viewing_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
325 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
326 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
327 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
328 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
329 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
330 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
331 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
332 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
333 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
334 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
335 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
336 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
337 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
338 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
339 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
340 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
341 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
342 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
343 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_19 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
344 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
345 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_21 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
346 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_22 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
347 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_23 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
348 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1235641 | Viewing: Reference lines and Scale annotators | NA | 4.3.8.Cardiac (MR) Viewer | 1418191 | Reference_Line_24 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
349 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352866 | All_series_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
350 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352866 | All_series_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
351 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352866 | All_series_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
352 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352866 | All_series_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
353 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352866 | All_series_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
354 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352866 | All_series_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
355 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
356 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
357 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
358 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
359 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
360 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
361 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
362 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
363 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
364 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
365 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352867 | Zoom_to_point_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
366 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352917 | TimeLine_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
367 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352917 | TimeLine_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
368 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352917 | TimeLine_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
369 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352917 | TimeLine_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
370 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352917 | TimeLine_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
371 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352917 | TimeLine_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
372 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352921 | Viewing_Page_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
373 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352921 | Viewing_Page_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
374 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352921 | Viewing_Page_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
375 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352921 | Viewing_Page_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
376 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352921 | Viewing_Page_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
377 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352921 | Viewing_Page_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
378 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352921 | Viewing_Page_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
379 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352921 | Viewing_Page_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
380 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352921 | Viewing_Page_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
381 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
382 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
383 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
384 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
385 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
386 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
387 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
388 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
389 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
390 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
391 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353103 | Single_series_scroll_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
392 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353112 | Planar_mode_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
393 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353112 | Planar_mode_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
394 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353112 | Planar_mode_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
395 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353112 | Planar_mode_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
396 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353112 | Planar_mode_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
397 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353112 | Planar_mode_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
398 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1353112 | Planar_mode_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
399 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1360497 | Top_slice_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
400 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1360497 | Top_slice_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
401 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1360497 | Top_slice_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
402 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1360497 | Top_slice_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
403 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1360497 | Top_slice_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
404 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1360497 | Top_slice_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
405 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241436 | User preferences: Hanging Protocol page | NA | 4.3.8.Cardiac (MR) Viewer | 1365127 | User_Preference_HP1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
406 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241436 | User preferences: Hanging Protocol page | NA | 4.3.8.Cardiac (MR) Viewer | 1365127 | User_Preference_HP2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
407 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241436 | User preferences: Hanging Protocol page | NA | 4.3.8.Cardiac (MR) Viewer | 1365127 | User_Preference_HP3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
408 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241436 | User preferences: Hanging Protocol page | NA | 4.3.8.Cardiac (MR) Viewer | 1365127 | User_Preference_HP4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
409 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241436 | User preferences: Hanging Protocol page | NA | 4.3.8.Cardiac (MR) Viewer | 1365127 | User_Preference_HP5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
410 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241436 | User preferences: Hanging Protocol page | NA | 4.3.8.Cardiac (MR) Viewer | 1365127 | User_Preference_HP6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
411 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241436 | User preferences: Hanging Protocol page | NA | 4.3.8.Cardiac (MR) Viewer | 1365127 | User_Preference_HP7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
412 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241436 | User preferences: Hanging Protocol page | NA | 4.3.8.Cardiac (MR) Viewer | 1365127 | User_Preference_HP8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
413 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241436 | User preferences: Hanging Protocol page | NA | 4.3.8.Cardiac (MR) Viewer | 1365127 | User_Preference_HP9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
414 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241436 | User preferences: Hanging Protocol page | NA | 4.3.8.Cardiac (MR) Viewer | 1365127 | User_Preference_HP10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
415 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
416 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
417 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
418 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
419 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
420 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
421 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
422 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_8 | PASS | NA |  | Pass
423 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
424 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
425 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
426 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
427 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
428 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
429 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
430 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
431 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
432 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
433 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_19 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
434 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
435 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_21 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
436 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_22 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
437 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_23 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
438 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_24 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
439 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_25 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
440 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_26 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
441 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
442 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
443 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
444 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
445 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
446 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
447 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
448 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
449 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
450 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
451 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
452 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
453 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
454 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
455 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
456 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
457 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
458 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
459 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
460 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
461 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
462 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
463 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
464 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
465 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
466 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
467 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
468 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
469 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241456 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | Preference_Common_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
470 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1351127 | User preferences: common behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1368837 | User_Preferences_Units setting_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
471 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
472 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
473 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
474 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
475 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
476 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
477 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
478 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
479 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
480 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
481 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241393 | User preferences: Viewing page | NA | 4.3.8.Cardiac (MR) Viewer | 1352870 | Auto_Reg_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
482 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
483 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
484 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
485 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
486 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
487 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
488 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
489 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
490 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
491 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
492 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
493 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
494 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
495 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362946 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
496 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362947 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
497 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362947 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
498 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362947 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
499 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240466 | Automatic capturing | NA | 4.3.8.Cardiac (MR) Viewer | 1362947 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
500 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
501 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
502 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
503 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
504 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
505 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
506 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
507 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
508 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
509 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
510 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
511 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
512 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
513 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
514 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
515 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
516 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
517 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
518 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_20 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
519 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_21 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
520 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_22 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
521 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_23 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
522 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_24 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
523 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1240464 | Manual capturing of images as new series | NA | 4.3.8.Cardiac (MR) Viewer | 1363466 | Step_25 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
524 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
525 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
526 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
527 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
528 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
529 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
530 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
531 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
532 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
533 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
534 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
535 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
536 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
537 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
538 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
539 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
540 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
541 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318009 | Operations on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1368817 | Step_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
542 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
543 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
544 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
545 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
546 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
547 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
548 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
549 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
550 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
551 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
552 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
553 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
554 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
555 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
556 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
557 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
558 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
559 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
560 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1474781 | Step_19 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
561 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318000 | Save Findgins | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
562 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
563 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
564 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
565 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
566 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
567 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
568 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
569 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
570 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1317993 | Findings types | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
571 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318004 | Quick Review on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
572 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318004 | Quick Review on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
573 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318004 | Quick Review on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
574 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318004 | Quick Review on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
575 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318004 | Quick Review on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
576 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318004 | Quick Review on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
577 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318004 | Quick Review on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
578 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318004 | Quick Review on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
579 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318004 | Quick Review on findings | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
580 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1413756 | Send to film (basic workflow) | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
581 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1413756 | Send to film (basic workflow) | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
582 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1413756 | Send to film (basic workflow) | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
583 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1413756 | Send to film (basic workflow) | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
584 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1413756 | Send to film (basic workflow) | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
585 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1413756 | Send to film (basic workflow) | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
586 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1413756 | Send to film (basic workflow) | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
587 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1413756 | Send to film (basic workflow) | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
588 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1413756 | Send to film (basic workflow) | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
589 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318000 | Save Findgins | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
590 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318000 | Save Findgins | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
591 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318000 | Save Findgins | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
592 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318000 | Save Findgins | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
593 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318000 | Save Findgins | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
594 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318000 | Save Findgins | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
595 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318000 | Save Findgins | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
596 | 1179491 | Results Export | 1182564 1182563 | Exporting findings Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1318000 | Save Findgins | NA | 4.3.8.Cardiac (MR) Viewer | 1475511 | Step_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
597 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1395433 | Save and Launch Results | NA | 4.3.8.Cardiac (MR) Viewer | 1490056 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
598 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1395433 | Save and Launch Results | NA | 4.3.8.Cardiac (MR) Viewer | 1490056 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
599 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1395433 | Save and Launch Results | NA | 4.3.8.Cardiac (MR) Viewer | 1490056 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
600 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1395433 | Save and Launch Results | NA | 4.3.8.Cardiac (MR) Viewer | 1490056 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
601 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1395433 | Save and Launch Results | NA | 4.3.8.Cardiac (MR) Viewer | 1490056 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
602 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1395433 | Save and Launch Results | NA | 4.3.8.Cardiac (MR) Viewer | 1490056 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
603 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1395433 | Save and Launch Results | NA | 4.3.8.Cardiac (MR) Viewer | 1490056 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
604 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1438347 | Result Summary card | NA | 4.3.8.Cardiac (MR) Viewer | 1490244 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
605 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1438347 | Result Summary card | NA | 4.3.8.Cardiac (MR) Viewer | 1490244 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
606 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1438347 | Result Summary card | NA | 4.3.8.Cardiac (MR) Viewer | 1490244 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
607 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1438347 | Result Summary card | NA | 4.3.8.Cardiac (MR) Viewer | 1490244 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
608 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1438347 | Result Summary card | NA | 4.3.8.Cardiac (MR) Viewer | 1490244 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
609 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1438347 | Result Summary card | NA | 4.3.8.Cardiac (MR) Viewer | 1490244 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
610 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1438347 | Result Summary card | NA | 4.3.8.Cardiac (MR) Viewer | 1490244 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
611 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
612 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
613 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
614 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
615 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
616 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
617 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
618 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
619 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
620 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
621 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
622 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
623 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
624 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
625 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
626 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
627 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
628 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
629 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_19 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
630 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
631 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_21 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
632 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_22 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
633 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_23 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
634 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_24 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
635 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_25 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
636 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_26 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
637 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_27 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
638 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_28 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
639 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_29 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
640 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_30 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
641 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_31 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
642 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_32 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
643 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_33 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
644 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_34 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
645 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_35 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
646 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_36 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
647 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_37 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
648 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_38 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
649 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_39 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
650 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_40 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
651 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_41 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
652 | 1179490 | Suite Workflow | 1182563 | Finding dashboard (managment) | Section 2.2.1 and Section 2.2.2 | 1439304 | Insights (Rule based notification) | NA | 4.3.8.Cardiac (MR) Viewer | 1490245 | Step_42 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
653 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
654 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
655 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
656 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
657 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
658 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
659 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
660 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
661 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
662 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
663 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
664 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_19 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
665 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
666 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_21 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
667 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_22 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
668 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_23 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
669 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_24 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
670 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_25 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
671 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241417 | User preferences: Keyboard shortcuts page | NA | 4.3.8.Cardiac (MR) Viewer | 1365148 | Keyboard_shortcut_26 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
672 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
673 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
674 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
675 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
676 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
677 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
678 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
679 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
680 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
681 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
682 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
683 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
684 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1241429 | User preferences: Mouse + Key Manipulation page | NA | 4.3.8.Cardiac (MR) Viewer | 1365206 | Mouse_Key_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
685 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
686 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
687 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
688 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
689 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
690 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
691 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
692 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
693 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
694 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
695 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
696 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
697 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
698 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
699 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
700 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
701 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238789 | Viewing: Slab mode | NA | 4.3.8.Cardiac (MR) Viewer | 1418149 | Viewing_mode_17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
702 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
703 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240893 | HP: Private Vs. Public HPs | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
704 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
705 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
706 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
707 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
708 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240893 | HP: Private Vs. Public HPs | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
709 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
710 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
711 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240893 | HP: Private Vs. Public HPs | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
712 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
713 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
714 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240893 | HP: Private Vs. Public HPs | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
715 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
716 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240893 | HP: Private Vs. Public HPs | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
717 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
718 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240893 | HP: Private Vs. Public HPs | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
719 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
720 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240913 | HP: Activating HP | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
721 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
722 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240893 | HP: Private Vs. Public HPs | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
723 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
724 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240917 | HP Logic: Defining default HP per user | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
725 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240891 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
726 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240917 | HP: Creating new hanging protocol | NA | 4.3.8.Cardiac (MR) Viewer | 1360491 | HP_16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
727 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240918 | Viewer: HP: Hanging protocols containing multiple exams | NA | 4.3.8.Cardiac (MR) Viewer | 1360506 | Multi-Study_HP_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
728 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240918 | Viewer: HP: Hanging protocols containing multiple exams | NA | 4.3.8.Cardiac (MR) Viewer | 1360506 | Multi-Study_HP_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
729 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240918 | Viewer: HP: Hanging protocols containing multiple exams | NA | 4.3.8.Cardiac (MR) Viewer | 1360506 | Multi-Study_HP_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
730 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240918 | Viewer: HP: Hanging protocols containing multiple exams | NA | 4.3.8.Cardiac (MR) Viewer | 1360506 | Multi-Study_HP_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
731 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240918 | Viewer: HP: Hanging protocols containing multiple exams | NA | 4.3.8.Cardiac (MR) Viewer | 1360506 | Multi-Study_HP_5 | Fail | 1499697 | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Defect no paln to fix
732 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240918 | Viewer: HP: Hanging protocols containing multiple exams | NA | 4.3.8.Cardiac (MR) Viewer | 1360506 | Multi-Study_HP_6 | Fail | 1499697 | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Defect no paln to fix
733 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240918 | Viewer: HP: Hanging protocols containing multiple exams | NA | 4.3.8.Cardiac (MR) Viewer | 1360506 | Multi-Study_HP_7 | NA | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Dependent on step 'Multi-Study_HP_6' hence cannot be executed.
734 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240918 | Viewer: HP: Hanging protocols containing multiple exams | NA | 4.3.8.Cardiac (MR) Viewer | 1360506 | Multi-Study_HP_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
735 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240918 | Viewer: HP: Hanging protocols containing multiple exams | NA | 4.3.8.Cardiac (MR) Viewer | 1360506 | Multi-Study_HP_9 | Fail | 1499697 | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Defect no paln to fix
736 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240918 | Viewer: HP: Hanging protocols containing multiple exams | NA | 4.3.8.Cardiac (MR) Viewer | 1360506 | Multi-Study_HP_10 | NA | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Dependent on step 'Multi-Study_HP_9' hence cannot be executed.
737 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240929 | Viewer: Factory hanging protocol (based on the loaded data) | NA | 4.3.8.Cardiac (MR) Viewer | 1360510 | Factory_hanging_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
738 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240929 | Viewer: Factory hanging protocol (based on the loaded data) | NA | 4.3.8.Cardiac (MR) Viewer | 1360510 | Factory_hanging_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
739 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240929 | Viewer: Factory hanging protocol (based on the loaded data) | NA | 4.3.8.Cardiac (MR) Viewer | 1360510 | Factory_hanging_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
740 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240929 | Viewer: Factory hanging protocol (based on the loaded data) | NA | 4.3.8.Cardiac (MR) Viewer | 1360510 | Factory_hanging_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
741 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240929 | Viewer: Factory hanging protocol (based on the loaded data) | NA | 4.3.8.Cardiac (MR) Viewer | 1360510 | Factory_hanging_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
742 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240929 | Viewer: Factory hanging protocol (based on the loaded data) | NA | 4.3.8.Cardiac (MR) Viewer | 1360510 | Factory_hanging_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
743 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240929 | Viewer: Factory hanging protocol (based on the loaded data) | NA | 4.3.8.Cardiac (MR) Viewer | 1360510 | Factory_hanging_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
744 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240932 | HP: HP list tools and behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1363360 | HP_list_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
745 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240932 | HP: HP list tools and behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1363360 | HP_list_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
746 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240932 | HP: HP list tools and behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1363360 | HP_list_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
747 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240932 | HP: HP list tools and behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1363360 | HP_list_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
748 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240932 | HP: HP list tools and behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1363360 | HP_list_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
749 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240932 | HP: HP list tools and behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1363360 | HP_list_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
750 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240932 | HP: HP list tools and behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1363360 | HP_list_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
751 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240932 | HP: HP list tools and behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1363360 | HP_list_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
752 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240932 | HP: HP list tools and behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1363360 | HP_list_9 | Fail | 1499697 | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Defect no paln to fix
753 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240932 | HP: HP list tools and behavior | NA | 4.3.8.Cardiac (MR) Viewer | 1363360 | HP_list_10 | NA | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | This test step is dependent on step HP_list_9 hence cannot be executed.
754 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240934 | HP: import HP from other users | NA | 4.3.8.Cardiac (MR) Viewer | 1364264 | Import_HP_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
755 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240934 | HP: import HP from other users | NA | 4.3.8.Cardiac (MR) Viewer | 1364264 | Import_HP_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
756 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240934 | HP: import HP from other users | NA | 4.3.8.Cardiac (MR) Viewer | 1364264 | Import_HP_3 | Fail | 1499126 | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Defect no paln to fix
757 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240934 | HP: import HP from other users | NA | 4.3.8.Cardiac (MR) Viewer | 1364264 | Import_HP_4 | NA | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Test step is dependent on step Import_HP_3 hence cannot be executed.
758 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240934 | HP: import HP from other users | NA | 4.3.8.Cardiac (MR) Viewer | 1364264 | Import_HP_5 | NA | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Test step is dependent on step Import_HP_3 hence cannot be executed.
759 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1240934 | HP: import HP from other users | NA | 4.3.8.Cardiac (MR) Viewer | 1364264 | Import_HP_6 | NA | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Test step is dependent on step Import_HP_3 hence cannot be executed.
760 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360647 | HP_Labelling_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
761 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360647 | HP_Labelling_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
762 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360647 | HP_Labelling_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
763 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360647 | HP_Labelling_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
764 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360647 | HP_Labelling_5 | NA | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Label data is supported only for MR Modality hence testcase needs to be removed.
765 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360647 | HP_Labelling_6 | NA | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | Label data is supported only for MR Modality hence testcase needs to be removed.
766 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360648 | HP_Labelling_Config_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
767 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360648 | HP_Labelling_Config_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
768 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360648 | HP_Labelling_Config_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
769 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360648 | HP_Labelling_Config_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
770 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.8.Cardiac (MR) Viewer | 1360648 | HP_Labelling_Config_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
771 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
772 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
773 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
774 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
775 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
776 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
777 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
778 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
779 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
780 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
781 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
782 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
783 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
784 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
785 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
786 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
787 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
788 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step18 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
789 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step19 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
790 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step20 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
791 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step21 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
792 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step22 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
793 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step23 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
794 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step24 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
795 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step25 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
796 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step26 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
797 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step27 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
798 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step28 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
799 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step29 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
800 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step30 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
801 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step31 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
802 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step32 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
803 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step33 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
804 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step34 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
805 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step35 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
806 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step36 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
807 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step37 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
808 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step38 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
809 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step39 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
810 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step40 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
811 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step41 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
812 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.8.Cardiac (MR) Viewer | 1362945 | Step42 | PASS | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
813 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2

## Traceability Matrix Func LVRV
DESIGN INPUTS | [10 empty] | DESIGN VERIFICATION |  |  | DESIGN VALIDATION
SL No | URS Requirement ID | URS Feature (DHF380082,LH) | PRS Requirement ID | PRS Feature (DHF380211,Rev LJ ) | SDS ID (DHF380234 ,Rev LE) | DRS Requirement ID | DRS Feature (D000856146,Rev I) | RMM ID (D000879224 , Rev F) | DDS Id (D000993081,Rev D) | Verification Test Case (ID / Brief Description) (D000996525 ,Rev I) | Verification Test Step ID (D000996525 ,Rev I) | Verification Test Result (Pass/Fail) [Verification_Report: D000996308, Rev E] Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | Defect ID (for Failed steps from Verification) | Validation Test ID (Validation Plan DHF380226, Rev D, Validation Report D000926820 , Rev C) | Validation Test Result (Pass/Fail) [Validation Report D000926820, Rev C] | Defect ID (for Failed steps from Validation) | Remarks
1 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
2 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
3 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
4 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
5 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
6 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
7 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
8 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
9 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
10 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
11 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
12 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366589 | Data_Validation_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
13 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
14 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
15 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
16 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
17 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
18 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
19 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
20 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
21 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
22 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
23 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
24 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
25 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
26 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
27 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.3 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
28 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.4 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
29 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.5 | 1395181 | ED/ES - initial location | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
30 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.5 | 1395416 | Apply automatic segmentation manually | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
31 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.6 | 1395456 | Save results | NA | 4.3.2 Functional Volumetric Application | 1366630 | Preprocess_9 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
32 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
33 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
34 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1395076 | LV and RV analysis | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
35 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1395181 | ED/ES - initial location | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
36 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1395076 | LV and RV analysis | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
37 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285032 | Set ED/ES - user operations | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
38 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1395076 | LV and RV analysis | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
39 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285032 | Set ED/ES - user operations | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
40 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1395076 | LV and RV analysis | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
41 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
42 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1395181 | ED/ES - initial location | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
43 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1240463 | Attributes to be saved with the finding | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
44 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1240463 | Attributes to be saved with the finding | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
45 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1240463 | Attributes to be saved with the finding | NA | 4.3.2 Functional Volumetric Application | 1366682 | LVRV_Common_WF_9 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
46 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
47 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
48 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
49 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1395087 | Task guidance and Single stage application | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
50 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1395087 | Task guidance and Single stage application | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
51 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285083 | Edit contour options and tools | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
52 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285109 | Propagate edited contours | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
53 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1412553 | Edit contour - undo | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
54 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285088 | Delete contours | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
55 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285087 | Create a new contour | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
56 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285088 | Delete contours | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
57 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285087 | Create a new contour | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
58 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285088 | Delete contours | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
59 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285087 | Create a new contour | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
60 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285088 | Delete contours | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_9 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
61 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285085 | Recalculate LV contours (semi automatic algorithm) | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_9 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
62 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285113 | Spokes wheel | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_10 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
63 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1240463 | Attributes to be saved with the finding | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_11 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
64 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1240463 | Attributes to be saved with the finding | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_12 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
65 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1240463 | Attributes to be saved with the finding | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_13 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
66 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284991 | LV and RV analysis | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_14 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
67 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285046 | Set ED/ES hanging protocols | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_15 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
68 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_16 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
69 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_16 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
70 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1284988 | Supported Data | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_17 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
71 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285304 | LV and RV automatic segmentation | NA | 4.3.2 Functional Volumetric Application | 1366717 | LVRV_Seg_17 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
72 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285265 | Blood pool based measurements | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
73 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1233675 | Launching external applications | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
74 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285265 | Blood pool based measurements | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
75 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285297 | Result table | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
76 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285221 | Measurements | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
77 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285221 | Measurements | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
78 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285258 | Volume Graph | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
79 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285258 | Volume Graph | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
80 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1240463 | Attributes to be saved with the finding | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
81 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285087 | Create a new contour | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
82 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285089 | 3D editing options | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
83 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285083 | Edit contour options and tools | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
84 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285087 | Create a new contour | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
85 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285109 | Propagate edited contours | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
86 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285265 | Blood pool based measurements | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
87 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285221 | Measurements | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_9 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
88 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1293365 | Normal values | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_9 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
89 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285286 | Heart rate | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_10 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
90 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285294 | Patient details | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_11 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
91 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285240 | Result stage hanging protocols | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_12 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
92 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285294 | Patient details | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_12 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
93 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285221 | Measurements | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_12 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
94 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285235 | Polar maps | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_13 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
95 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285113 | Spokes wheel | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_13 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
96 | 1044734 | Filming | 529314 | ISPV7_PRS_ID_2_1: Film & Report | Section 2.2.1 and Section 2.2.3 | 1410882 | Save as option for images | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_14 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
97 | 1044734 | Filming | 529314 | ISPV7_PRS_ID_2_1: Film & Report | Section 2.2.1 and Section 2.2.4 | 1413756 | Send to film (basic workflow) | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_15 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
98 | 688416 | ISP [URS] | 529698 | ISPV7_PRS_ID_7_06: Report Integration with RIS and EMR Systems | Section 2.2.1 and Section 2.2.5 | 1440092 | Copy Table to clipboard | NA | 4.3.2 Functional Volumetric Application | 1366718 | LVRV_Result_16 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
99 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
100 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
101 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
102 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
103 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
104 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
105 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
106 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
107 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
108 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
109 | 1179490 | Suite Workflow | 1182550 | Multiple applications side by side | Section 2.2.1 and Section 2.2.2 | 1288859 | Opening new applications | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
110 | 1179490 | Suite Workflow | 1182550 | Multiple applications side by side | Section 2.2.1 and Section 2.2.2 | 1233673 | Multiple Applications sides by side | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
111 | 1179490 | Suite Workflow | 1182550 | Multiple applications side by side | Section 2.2.1 and Section 2.2.2 | 1288854 | Displaying multiple applications side by side | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
112 | 1179490 | Suite Workflow | 1182550 | Multiple applications side by side | Section 2.2.1 and Section 2.2.2 | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
113 | 1179457 | Conversion to semantic model | 1436922 | Discrete Reporting | Section 2.2.1 and Section 2.2.3 | 1440093 | Discrete Reporting | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367960 | CMRLVRV_Workflow4_8 | Pass | NA | DHF380226>12.4 | VbA | VbA | None
114 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.4 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_1 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
115 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.5 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_2 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
116 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.6 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_3 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
117 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.7 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_4 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
118 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.8 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_5 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
119 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.9 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_6 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
120 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.10 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_7 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
121 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.11 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_8 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
122 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.12 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_9 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
123 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.13 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_10 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
124 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.14 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_11 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
125 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.15 | 1410921 | User preferences: Normal Values Editor | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1441320 | CMRLVRV_NormalValueEditor_12 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
126 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300209 | fCMR: Review/Edit contours - Usability safety mitigations | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
127 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1360769 | fCMR : Notify user on automatic results | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
128 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1360769 | fCMR : Notify user on automatic results | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
129 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300209 | fCMR: Review/Edit contours - Usability safety mitigations | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
130 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1360769 | fCMR : Notify user on automatic results | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
131 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1360769 | fCMR : Notify user on automatic results | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
132 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300206 | fCMR: Set ED/ES - Usability safety mitigations | FCMR.WrongED/ES.UserAwarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
133 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300207 | fCMR: Change ED/ES in after edit contours - Usability safety mitigations | FCMR.EditingED/ES.UserAwarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
134 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300209 | fCMR: Review/Edit contours - Usability safety mitigations | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
135 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1360769 | fCMR : Notify user on automatic results | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
136 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1360769 | fCMR : Notify user on automatic results | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
137 | 1182563 | Displaying findings from all applications | 1360769 | fCMR : Notify user on automatic results | Section 2.2.1 and Section 2.2.2 | 1360255 | fCMR : Showing reference lines | FCMR.Referenceline | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
138 | 1228262 | General Behavior | 1360769 | fCMR : Notify user on automatic results | Section 2.2.1 and Section 2.2.3 | 1360255 | fCMR : Showing reference lines | FCMR.Referenceline | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
139 | 1182563 | Displaying findings from all applications | 1360769 | fCMR : Notify user on automatic results | Section 2.2.1 and Section 2.2.2 | 1360255 | fCMR : Showing reference lines | FCMR.Referenceline | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
140 | 1228262 | General Behavior | 1360769 | fCMR : Notify user on automatic results | Section 2.2.1 and Section 2.2.3 | 1360255 | fCMR : Showing reference lines | FCMR.Referenceline | 4.3.2 Functional Volumetric Application | 1366742 | LVRV_Haz_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
141 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | FCMR.IncorrectSpokeWheel.Userawarness | 4.3.1 Cardiac.Common | 1366742 | LVRV_Haz_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
142 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | FCMR.PolarMapsWrongValues | 4.3.1 Cardiac.Common | 1366742 | LVRV_Haz_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
143 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | FMCR.PolarMapsWallMotion.Userawarness | 4.3.1 Cardiac.Common | 1366742 | LVRV_Haz_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
144 | 1179457 | Conversion to semantic model | 1436922 | Discrete Reporting | Section 2.2.1 and Section 2.2.2 | 1440093 | Discrete Reporting | FCMR.DiscreteReporting | 4.3.1 Cardiac.Common | 1366742 | LVRV_Haz_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
145 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300209 | fCMR: Review/Edit contours - Usability safety mitigations | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
146 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1360769 | fCMR : Notify user on automatic results | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
147 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1360769 | fCMR : Notify user on automatic results | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
148 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300209 | fCMR: Review/Edit contours - Usability safety mitigations | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
149 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1360769 | fCMR : Notify user on automatic results | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
150 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1360769 | fCMR : Notify user on automatic results | FCMR.IncorrectAuto Contour.Userawarness | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
151 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | FCMR.IncorrectSpokeWheel.Userawarness | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
152 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | FCMR.PolarMapsWrongValues | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
153 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | FMCR.PolarMaps.Userawarness | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
154 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | FMCR.PolarMapsWallMotion.Userawarness | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
155 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300210 | fCMR: Review results - Usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
156 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
157 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300210 | fCMR: Review results - Usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
158 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
159 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.2.1.1 Cardiac Suite (UI) | 1366745 | LVRV_Haz_Results_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
160 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300210 | fCMR: Review results - Usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
161 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
162 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.2.1.1 Cardiac Suite (UI) | 1366745 | LVRV_Haz_Results_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
163 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300212 | fCMR: Change calculation method (Endo vbs Blood pool) | FCMR.BloodPool.Userawarness | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
164 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300210 | fCMR: Review results - Usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
165 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
166 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | FCMR.TabelMeasurement.Userawarness. | 4.2.1.1 Cardiac Suite (UI) | 1366745 | LVRV_Haz_Results_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
167 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300210 | fCMR: Review results - Usability safety mitigations | FCMR.BSAValues.Userawarness. | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
168 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | FCMR.BSAValues.Userawarness. | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
169 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | FCMR.BSAValues.Userawarness. | 4.2.1.1 Cardiac Suite (UI) | 1366745 | LVRV_Haz_Results_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
170 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300211 | fCMR: Volume Graph - usability safety mitigations | FMCR.Graph.Userawarness | 4.3.2 Functional Volumetric Application | 1366745 | LVRV_Haz_Results_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
171 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300186 | Review graph - usability safety mitigations | FMCR.Graph.Userawarness | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
172 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1300188 | Patient or Exam related parameters (user editable) - usability safety mitigations | FCMR.BSAValues.Userawarness. | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
173 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1300189 | Send results to finding dashboard - usability safety mitigations | FCMR.FindingsDashResultUpdate.Userawarness. | 4.3.1 Cardiac.Common | 1366745 | LVRV_Haz_Results_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
174 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355342 | Common: Incorrect display of Pixel data due to 2D image manipulations | 4.3.2 | 4.1 System Architecture | 1367569 | Common_RMM_01 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
175 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355352 | Common: The user is scrolling images in a series using mouse click and drag. The user is scrolling in a speed that causes the system to skip slices. The user is unaware that Not all images have been shown | 4.3.12 | 4.1 System Architecture | 1367569 | Common_RMM_01 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
176 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355344 | Common: Wrong or missing contrast indication on displayed image | 4.3.4 | 4.1 System Architecture | 1367569 | Common_RMM_02 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
177 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355353 | Common: User loaded 2 series with different field of view. The “zoom link” is causing a part of the anatomy to be out of view The user think he is looking at the image in the original zoom | 4.3.14 | 4.1 System Architecture | 1367569 | Common_RMM_03 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
178 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355354 | Common: Showing wrong slice number under "phase" annotation in addition to the correct slice number under the "slice" annotation | 4.3.15 | 4.1 System Architecture | 1367569 | Common_RMM_04 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
179 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1359246 | Common: Viewing or Analysis based on incorrect or wrong merging of split series | 4.10. | 4.1 System Architecture | 1367569 | Common_RMM_05 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
180 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1356011 | Common: Geometrical measurement value or position is changed after activating a saved finding[4.4.7 ] | 4.4.7 | 4.1 System Architecture | 1366755 | MRC_RMF_1 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
181 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1356011 | Common: Geometrical measurement value or position is changed after activating a saved finding[4.4.7 ] | 4.4.7 | 4.1 System Architecture | 1366755 | MRC_RMF_1 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
182 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 4.4.8 | 4.1 System Architecture | 1366755 | MRC_RMF_2 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
183 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 4.4.8 | 4.1 System Architecture | 1366755 | MRC_RMF_2 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
184 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357606 | Common: Incorrect or incorrect presentation of scale, grid on images.[ 4.4.15] | 4.4.15 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_1 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
185 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357610 | Common: Textual annotation's position is changed due to image manipulations (e.g. zoom, pan, sweivel, rotate, flip)[ 4.4.20] | 4.4.20 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_2 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
186 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357620 | Common: Wrong interpretation of images when lossy compression is used 2. Wrong interpretation of images when lossy compression with resolution reduction is used[ 4.5.1] | 4.5.1 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_3 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
187 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357871 | Common: new images are saved to a wrong exam of the same patient | 4.7.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_1 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
188 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
189 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.2 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
190 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
191 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1, | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
192 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.2 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
193 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
194 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1, | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
195 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.2 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
196 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
197 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357897 | Common: Wrong correlation between images. Incorrect position of location represented in one or more of the viewports. | 4.11 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_5 | Pass | NA | DHF380226 >12.4 D000926820>3.3.8 | Pass | None | None
198 | 1179486 | Viewer application | 1182554 926498 | Viewer ISPV11_PRS_ID_1_2: Dual Monitor Support | Section 2.2.1 and Section 2.2.2 | DRS: 1235590 Defect: 1461893 | DRS: Viewing: Fit to monitor Defect: MRC Suite launch in Single Phase-Dual Monitor display | NA | 4.3.8.Cardiac (MR) Viewer | 1479409 | DualMonitor1 | Pass | NA | NA | Pass | None | None
199 | 1179486 | Viewer application | 1182554 926498 | Viewer ISPV11_PRS_ID_1_2: Dual Monitor Support | Section 2.2.1 and Section 2.2.2 | DRS: 1235590 Defect: 1461893 | DRS: Viewing: Fit to monitor Defect: MRC Suite launch in Single Phase-Dual Monitor display | NA | 4.3.8.Cardiac (MR) Viewer | 1479409 | DualMonitor2 | Pass | NA | NA | Pass | None | None
200 | 1179486 | Viewer application | 1182554 926498 | Viewer ISPV11_PRS_ID_1_2: Dual Monitor Support | Section 2.2.1 and Section 2.2.2 | DRS: 1235590 Defect: 1461893 | DRS: Viewing: Fit to monitor Defect: MRC Suite launch in Single Phase-Dual Monitor display | NA | 4.3.8.Cardiac (MR) Viewer | 1479409 | DualMonitor3 | Pass | NA | NA | Pass | None | None
201 | 1179486 | Viewer application | 1182554 926498 | Viewer ISPV11_PRS_ID_1_2: Dual Monitor Support | Section 2.2.1 and Section 2.2.2 | DRS: 1235590 Defect: 1461893 | DRS: Viewing: Fit to monitor Defect: MRC Suite launch in Single Phase-Dual Monitor display | NA | 4.3.8.Cardiac (MR) Viewer | 1479409 | DualMonitor4 | Pass | NA | NA | Pass | None | None
202 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | DRS: 1440092 Defect: 1479083 | Copy Table to clipboard | NA | 4.3.6. Temporal Enhancement Application | 1479416 | CopyClipboard1 | Pass | NA | NA | Pass | None | None
203 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | DRS: 1440092 Defect: 1479083 | Copy Table to clipboard | NA | 4.3.6. Temporal Enhancement Application | 1479416 | CopyClipboard2 | Pass | NA | NA | Pass | None | None
204 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | DRS: 1440092 Defect: 1479083 | Copy Table to clipboard | NA | 4.3.6. Temporal Enhancement Application | 1479416 | CopyClipboard3 | Pass | NA | NA | Pass | None | None
205 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | DRS: 1440092 Defect: 1479083 | Copy Table to clipboard | NA | 4.3.6. Temporal Enhancement Application | 1479416 | CopyClipboard4 | Pass | NA | NA | Pass | None | None
206 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328596 Defect: 1478383 | DRS: Segmentation & Result stage default hanging protocol and image position. Defect: T1 Mapping in MR Cardiac Analysis | N/A | 4.3.7 Mapping Application | 1479447 | T1Map1 | Pass | NA | NA | Pass | None | None
207 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328596 Defect: 1478383 | DRS: Segmentation & Result stage default hanging protocol and image position. Defect: T1 Mapping in MR Cardiac Analysis | N/A | 4.3.7 Mapping Application | 1479447 | T1Map2 | Pass | NA | NA | Pass | None | None
208 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328596 Defect: 1478383 | DRS: Segmentation & Result stage default hanging protocol and image position. Defect: T1 Mapping in MR Cardiac Analysis | N/A | 4.3.7 Mapping Application | 1479447 | T1Map3 | Pass | NA | NA | Pass | None | None
209 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328596 Defect: 1478383 | DRS: Segmentation & Result stage default hanging protocol and image position. Defect: T1 Mapping in MR Cardiac Analysis | N/A | 4.3.7 Mapping Application | 1479447 | T1Map4 | Pass | NA | NA | Pass | None | None
210 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328596 Defect: 1478383 | DRS: Segmentation & Result stage default hanging protocol and image position. Defect: T1 Mapping in MR Cardiac Analysis | N/A | 4.3.7 Mapping Application | 1479447 | T1Map5 | Pass | NA | NA | Pass | None | None
211 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328596 Defect: 1478383 | DRS: Segmentation & Result stage default hanging protocol and image position. Defect: T1 Mapping in MR Cardiac Analysis | N/A | 4.3.7 Mapping Application | 1479447 | T1Map6 | Pass | NA | NA | Pass | None | None
212 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328596 Defect: 1478383 | DRS: Segmentation & Result stage default hanging protocol and image position. Defect: T1 Mapping in MR Cardiac Analysis | N/A | 4.3.7 Mapping Application | 1479447 | T1Map7 | Pass | NA | NA | Pass | None | None
213 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328596 Defect: 1478383 | DRS: Segmentation & Result stage default hanging protocol and image position. Defect: T1 Mapping in MR Cardiac Analysis | N/A | 4.3.7 Mapping Application | 1479447 | T1Map8 | Pass | NA | NA | Pass | None | None
214 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | DRS: 1239033 Defect: 1429452 | DRS: Cine Widget Defect: Cine Linking across Viewports | NA | 4.3.8.Cardiac (MR) Viewer | 1479491 | CineLink_1 | Pass | NA | NA | Pass | None | None
215 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | DRS: 1239033 Defect: 1429452 | DRS: Cine Widget Defect: Cine Linking across Viewports | NA | 4.3.8.Cardiac (MR) Viewer | 1479491 | CineLink_2 | Pass | NA | NA | Pass | None | None
216 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | DRS: 1239033 Defect: 1429452 | DRS: Cine Widget Defect: Cine Linking across Viewports | NA | 4.3.8.Cardiac (MR) Viewer | 1479491 | CineLink_3 | Pass | NA | NA | Pass | None | None
217 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | DRS: 1285304 Defect: 1470103 | DRS: LV and RV automatic segmentation Defect: Error while loading | NA | 4.3.2 Functional Volumetric Application | 1479492 | Connection1 | Pass | NA | NA | Pass | None | None
218 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328579 Defect: 1429452 | DRS: Application launch Defect: Error while loading | N/A | 4.3.7 Mapping Application | 1479492 | Connection1 | Pass | NA | NA | Pass | None | None
219 | 135689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | DRS: 1410562 Defect: 1429452 | DRS: Loading the application Defect: Error while loading | NA | 4.3.6. Temporal Enhancement Application | 1479492 | Connection1 | Pass | NA | NA | Pass | None | None
220 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | DRS: 1285304 Defect: 1470103 | DRS: LV and RV automatic segmentation Defect: Error while loading | NA | 4.3.2 Functional Volumetric Application | 1479492 | Connection2 | Pass | NA | NA | Pass | None | None
221 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328579 Defect: 1429452 | DRS: Application launch Defect: Error while loading | N/A | 4.3.7 Mapping Application | 1479492 | Connection2 | Pass | NA | NA | Pass | None | None
222 | 135689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | DRS: 1410562 Defect: 1429452 | DRS: Loading the application Defect: Error while loading | NA | 4.3.6. Temporal Enhancement Application | 1479492 | Connection2 | Pass | NA | NA | Pass | None | None
223 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | DRS: 1285304 Defect: 1470103 | DRS: LV and RV automatic segmentation Defect: Error while loading | NA | 4.3.2 Functional Volumetric Application | 1479492 | Connection3 | Pass | NA | NA | Pass | None | None
224 | 1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | DRS: 1328579 Defect: 1429452 | DRS: Application launch Defect: Error while loading | N/A | 4.3.7 Mapping Application | 1479492 | Connection3 | Pass | NA | NA | Pass | None | None
225 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | DRS: 1410562 Defect: 1429452 | DRS: Loading the application Defect: Error while loading | NA | 4.3.6. Temporal Enhancement Application | 1479492 | Connection3 | Pass | NA | NA | Pass | None | None
226 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556

## Traceability Matrix Func LA LV
DESIGN INPUTS | [10 empty] | DESIGN VERIFICATION |  |  | DESIGN VALIDATION
SL No | URS Requirement ID | URS Feature (DHF380082,LH) | PRS Requirement ID | PRS Feature (DHF380211,Rev LJ ) | SDS ID (DHF380234 ,Rev LE) | DRS Requirement ID | DRS Feature (D000856146,Rev I) | RMM ID (D000879224 , Rev F) | DDS Id (D000993081,Rev D) | Verification Test Case (ID / Brief Description) (D000996542 ,Rev E) | Verification Test Step ID (D000996542 ,Rev E) | Verification Test Result (Pass/Fail) [Verification_Report: D000996308, Rev E] Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | Defect ID (for Failed steps from Verification) | Validation Test ID (Validation Plan DHF380226, Rev D, Validation Report D000926820 , Rev C) | Validation Test Result (Pass/Fail) [Validation Report D000926820, Rev C] | Defect ID (for Failed steps from Validation) | Remarks
1 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328899 | Data types | NA | 4.3.3 Functional Planar Application | 1367496 | FunLA_Val1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
2 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | D000993082: 9.3 Semantic Labeling 4.3.3 Functional Planar Application | 1367496 | FunLA_Val1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
3 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328899 | Data types | NA | 4.3.3 Functional Planar Application | 1367496 | FunLA_Val2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
4 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328899 | Data types | NA | 4.3.3 Functional Planar Application | 1367496 | FunLA_Val3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
5 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328899 | Data types | NA | 4.3.3 Functional Planar Application | 1367496 | FunLA_Val4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
6 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328899 | Data types | NA | 4.3.3 Functional Planar Application | 1367496 | FunLA_Val5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
7 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328899 | Data types | NA | 4.3.3 Functional Planar Application | 1367496 | FunLA_Val6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
8 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328899 | Data types | NA | 4.3.3 Functional Planar Application | 1367496 | FunLA_Val7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
9 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328899 | Data types | NA | 4.3.3 Functional Planar Application | 1367496 | FunLA_Val8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
10 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328899 | Data types | NA | 4.3.3 Functional Planar Application | 1367496 | FunLA_Val9 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
11 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328904 | Set ED/ES | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_1 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
12 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1332649 | Application workflow and hangings | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
13 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328904 | Set ED/ES | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_3 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
14 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1332649 | Application workflow and hangings | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_3 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
15 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328949 | LV segmentation | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_4 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
16 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328951 | Edit segmentation and midline | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_5 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
17 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328956 | Results | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_6 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
18 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1332649 | Application workflow and hangings | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_6 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
19 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328895 | General | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_6 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
20 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328956 | Results | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_7 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
21 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1358644 | Application results as finding | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_8 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
22 | 688416 | ISP [URS] | 529698 | ISPV7_PRS_ID_7_06: Report Integration with RIS and EMR Systems | Section 2.2.1 and Section 2.2.5 | 1440092 | Copy Table to clipboard | NA | 4.3.3 Functional Planar Application | 1367499 | FuncLA_WF1_9 | Pass | NA | NA | Pass | None | None
23 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328904 | Set ED/ES | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
24 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1332649 | Application workflow and hangings | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
25 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328904 | Set ED/ES | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
26 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1332649 | Application workflow and hangings | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
27 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328949 | LV segmentation | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
28 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328951 | Edit segmentation and midline | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
29 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328956 | Results | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
30 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1332649 | Application workflow and hangings | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
31 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328895 | General | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
32 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1328956 | Results | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
33 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1358644 | Application results as finding | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
34 | 1179457 | Conversion to semantic model | 1436922 | Discrete Reporting | Section 2.2.1 and Section 2.2.2 | 1440093 | Discrete Reporting | NA | 4.3.3 Functional Planar Application | 1367524 | FuncLA_WF2_9 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
35 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1234276 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
36 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
37 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1234276 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
38 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
39 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1234276 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
40 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
41 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1234276 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
42 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
43 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1288859 | Opening new applications | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
44 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1233673 | Multiple Applications sides by side | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
45 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1288854 | Displaying multiple applications side by side | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
46 | 1044734 | Filming | 529314 | ISPV7_PRS_ID_2_1: Film & Report | Section 2.2.1 and Section 2.2.3 | 1410882 | Save as option for images | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
47 | 1044734 | Filming | 529314 | ISPV7_PRS_ID_2_1: Film & Report | Section 2.2.1 and Section 2.2.4 | 1413756 | Send to film (basic workflow) | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1368838 | CMRLALV_Workflow3_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
48 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1361124 | Functional LA: Launched series are not perpendicular | FuncLA.AnalysisSeriesSelcetion.Userawarness | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
49 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1361131 | Functional LA: ED and ES phases setting | FuncLA.WrongDefaultED/ES.UserAwarness | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
50 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1361974 | Functional LA: Guide lines to th euser | FuncLA.Non-OptimalDefaultSlice.UserAwarness | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
51 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1361132 | Functional LA: Phases synch between series | FucnLA.ContoursOnIncorrectED/ES.Userawarness | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
52 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1350671 | Functional LA: Draw contours usability safety mitigations | FucnLA.MultiSliceWrongED/ES.Userawarness | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
53 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1361146 | Functional LA: Result table | FuncLA.TabelMeasurement.Userawarness. | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
54 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | FuncLA.TabelMeasurement.Userawarness. | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
55 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | FuncLA.TabelMeasurement.Userawarness. | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
56 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | FuncLA.TabelMeasurement.Userawarness. | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
57 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | FuncLA.BSAValues.Userawarness. | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
58 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1300188 | Patient or Exam related parameters (user editable) - usability safety mitigations | FuncLA.BSAValues.Userawarness | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
59 | 1353694 | MR Cardiac Long-Axis Functional Analysis application | 1182557 | Functional Long Axis | Section 2.2.1 and Section 2.2.2 | 1300189 | Send results to finding dashboard - usability safety mitigations | FuncLA.FindingsDashResultUpdate.Userawarness. | 4.3.3 Functional Planar Application | 1368292 | LALV_HAZ_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
60 | 1179457 | Conversion to semantic model | 1436922 | Discrete Reporting | Section 2.2.1 and Section 2.2.2 | 1440093 | Discrete Reporting | funcLA.DiscreteReporting | 4.3.1 Cardiac.Common | 1368292 | LALV_HAZ_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | NA | VbA | None | None
61 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355342 | Common: Incorrect display of Pixel data due to 2D image manipulations | 4.3.2 | 4.1 System Architecture | 1367569 | Common_RMM_01 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
62 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355352 | Common: The user is scrolling images in a series using mouse click and drag. The user is scrolling in a speed that causes the system to skip slices. The user is unaware that Not all images have been shown | 4.3.12 | 4.1 System Architecture | 1367569 | Common_RMM_01 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
63 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355344 | Common: Wrong or missing contrast indication on displayed image | 4.3.4 | 4.1 System Architecture | 1367569 | Common_RMM_02 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
64 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355353 | Common: User loaded 2 series with different field of view. The “zoom link” is causing a part of the anatomy to be out of view The user think he is looking at the image in the original zoom | 4.3.14 | 4.1 System Architecture | 1367569 | Common_RMM_03 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
65 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355354 | Common: Showing wrong slice number under "phase" annotation in addition to the correct slice number under the "slice" annotation | 4.3.15 | 4.1 System Architecture | 1367569 | Common_RMM_04 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
66 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1359246 | Common: Viewing or Analysis based on incorrect or wrong merging of split series | 4.10. | 4.1 System Architecture | 1367569 | Common_RMM_5 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
67 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1356011 | Common: Geometrical measurement value or position is changed after activating a saved finding[4.4.7 ] | 4.4.7 | 4.1 System Architecture | 1366755 | MRC_RMF_1 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
68 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1356011 | Common: Geometrical measurement value or position is changed after activating a saved finding[4.4.7 ] | 4.4.7 | 4.1 System Architecture | 1366755 | MRC_RMF_1 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
69 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 4.4.8 | 4.1 System Architecture | 1366755 | MRC_RMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
70 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 4.4.8 | 4.1 System Architecture | 1366755 | MRC_RMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
71 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357606 | Common: Incorrect or incorrect presentation of scale, grid on images.[ 4.4.15] | 4.4.15 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_1 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
72 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357610 | Common: Textual annotation's position is changed due to image manipulations (e.g. zoom, pan, sweivel, rotate, flip)[ 4.4.20] | 4.4.20 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
73 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357620 | Common: Wrong interpretation of images when lossy compression is used 2. Wrong interpretation of images when lossy compression with resolution reduction is used[ 4.5.1] | 4.5.1 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_3 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
74 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357871 | Common: new images are saved to a wrong exam of the same patient | 4.7.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_1 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
75 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
76 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.2 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
77 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
78 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1, | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
79 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.2 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
80 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
81 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1, | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
82 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.2 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
83 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None
84 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357897 | Common: Wrong correlation between images. Incorrect position of location represented in one or more of the viewports. | 4.11 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_5 | Pass | NA | DHF380226>12.4 D000926820>3.3.7 | Pass | None | None

## Traceability Matrix QFlow
DESIGN INPUTS | [10 empty] | DESIGN VERIFICATION |  |  | DESIGN VALIDATION
SL No | URS Requirement ID | URS Feature (DHF380082,LH) | PRS Requirement ID | PRS Feature (DHF380211,Rev LJ ) | SDS ID (DHF380234 ,Rev LE) | DRS Requirement ID | DRS Feature (D000856146,Rev I) | RMM ID (D000879224 , Rev F) | DDS Id (D000993081,Rev D) | Verification Test Case (ID / Brief Description) (D000996541 ,Rev E) | Verification Test Step ID (D000996541 ,Rev E) | Verification Test Result (Pass/Fail) [Verification_Report: D000996308, Rev E] Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | Defect ID (for Failed steps from Verification) | Validation Test ID (Validation Plan DHF380226, Rev D, Validation Report D000926820 , Rev C) | Validation Test Result (Pass/Fail) [Validation Report D000926820, Rev C] | Defect ID (for Failed steps from Validation) | Remarks
1 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | D000993082: 9.3 Semantic Labeling | 1367185 | CMRQflow_DV_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
2 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294383 | Supported data | NA | 4.3.5 QFlow Application | 1367185 | CMRQflow_DV_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
3 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294383 | Supported data | NA | 4.3.5 QFlow Application | 1367185 | CMRQflow_DV_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
4 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294383 | Supported data | NA | 4.3.5 QFlow Application | 1367185 | CMRQflow_DV_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
5 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294383 | Supported data | NA | 4.3.5 QFlow Application | 1367185 | CMRQflow_DV_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
6 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294383 | Supported data | NA | 4.3.5 QFlow Application | 1367185 | CMRQflow_DV_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
7 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294383 | Supported data | NA | 4.3.5 QFlow Application | 1367185 | CMRQflow_DV_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
8 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294384 | Viewports | NA | 4.3.5 QFlow Application | 1367187 | CMRQflow_Workflow1_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
9 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294385 | Slice phase matrix | NA | 4.3.5 QFlow Application | 1367187 | CMRQflow_Workflow1_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
10 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294390 | Cine Play | NA | 4.3.5 QFlow Application | 1367187 | CMRQflow_Workflow1_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
11 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294392 | Table settings | NA | 4.3.5 QFlow Application | 1367187 | CMRQflow_Workflow1_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
12 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294394 | Unit settings | NA | 4.3.5 QFlow Application | 1367187 | CMRQflow_Workflow1_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
13 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294402 | Velocity Probe | NA | 4.3.5 QFlow Application | 1367187 | CMRQflow_Workflow1_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
14 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294403 | Show overlay | NA | 4.3.5 QFlow Application | 1367187 | CMRQflow_Workflow1_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
15 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.3 | 1294386 | Show / Hide contours | NA | 4.3.5 QFlow Application | 1367187 | CMRQflow_Workflow1_8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
16 | 1044734 | Filming | 529314 | ISPV7_PRS_ID_2_1: Film & Report | Section 2.2.1 and Section 2.2.3 | 1410882 | Save as option for images | NA | 4.3.5 QFlow Application | 1367187 | CMRQflow_Workflow1_9 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
17 | 1044734 | Filming | 529314 | ISPV7_PRS_ID_2_1: Film & Report | Section 2.2.1 and Section 2.2.4 | 1413756 | Send to film (basic workflow) | NA | 4.3.5 QFlow Application | 1367187 | CMRQflow_Workflow1_10 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
18 | 688416 | ISP [URS] | 529698 | ISPV7_PRS_ID_7_06: Report Integration with RIS and EMR Systems | Section 2.2.1 and Section 2.2.5 | 1440092 | Copy Table to clipboard | NA | 4.3.2 Functional Volumetric Application | 1367187 | CMRQflow_Workflow1_11 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | NA | Pass | NA | None
19 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294435 | Create segmentation | NA | 4.3.5 QFlow Application | 1367190 | CMRQflow_Workflow2_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | VbA | None | None
20 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294436 | ROI labelling | NA | 4.3.5 QFlow Application | 1367190 | CMRQflow_Workflow2_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | VbA | None | None
21 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294438 | Contour verification and editing | NA | 4.3.5 QFlow Application | 1367190 | CMRQflow_Workflow2_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | VbA | None | None
22 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294436 | ROI labelling | NA | 4.3.5 QFlow Application | 1367190 | CMRQflow_Workflow2_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
23 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294504 | Delete ROI | NA | 4.3.5 QFlow Application | 1367190 | CMRQflow_Workflow2_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
24 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294506 | Pan ROI | NA | 4.3.5 QFlow Application | 1367190 | CMRQflow_Workflow2_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
25 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294508 | Rename contours | NA | 4.3.5 QFlow Application | 1367190 | CMRQflow_Workflow2_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
26 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294509 | Propagate contours | NA | 4.3.5 QFlow Application | 1367190 | CMRQflow_Workflow2_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
27 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294450 | E/A ratio for Mitral valve | NA | 4.3.5 QFlow Application | 1367190 | CMRQflow_Workflow2_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
28 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294456 | Temporal range | NA | 4.3.5 QFlow Application | 1367371 | CMRQflow_Workflow3_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
29 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294456 | Temporal range | NA | 4.3.5 QFlow Application | 1367371 | CMRQflow_Workflow3_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
30 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294464 | Background correction | NA | 4.3.5 QFlow Application | 1367371 | CMRQflow_Workflow3_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
31 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294464 | Background correction | NA | 4.3.5 QFlow Application | 1367371 | CMRQflow_Workflow3_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
32 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294478 | Results table | NA | 4.3.5 QFlow Application | 1367371 | CMRQflow_Workflow3_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
33 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294441 | Analysis Results table | NA | 4.3.5 QFlow Application | 1367371 | CMRQflow_Workflow3_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
34 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294416 | Set heart rate | NA | 4.3.5 QFlow Application | 1367371 | CMRQflow_Workflow3_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
35 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294442 | Range results table | NA | 4.3.5 QFlow Application | 1367371 | CMRQflow_Workflow3_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
36 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294443 | Current phase results table | NA | 4.3.5 QFlow Application | 1367371 | CMRQflow_Workflow3_8 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
37 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1294423 | BSA Indexed | NA | 4.3.5 QFlow Application | 1367371 | CMRQflow_Workflow3_9 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
38 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1300189 | Qflow.FindingsDashResultUpdate.Userawarness. | NA | 4.3.1 Cardiac.Common | 1367371 | CMRQflow_Workflow3_10 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
39 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1439158 | CSF flow direcion | NA | 4.3.1 Cardiac.Common | 1367371 | CMRQflow_Workflow3_11 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
40 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
41 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
42 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1234276 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
43 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
44 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
45 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1234276 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
46 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
47 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1234276 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
48 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
49 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
50 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_4 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
51 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
52 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_5 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
53 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1288859 | Opening new applications | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
54 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1233673 | Multiple Applications sides by side | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
55 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1288854 | Displaying multiple applications side by side | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_6 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
56 | 1179457 | Conversion to semantic model | 1436922 | Discrete Reporting | Section 2.2.1 and Section 2.2.2 | 1440093 | Discrete Reporting | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367608 | CMRQflow_Workflow4_7 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
57 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.3 | 1412026 | Qflow as a stand alone | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1456493 | CMRQflow_Standalone_1 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | NA | VbA | VbA | None
58 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.4 | 1412026 | Qflow as a stand alone | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1456493 | CMRQflow_Standalone_2 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | 1451555 | NA | VbA | VbA | None
59 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.5 | 1412026 | Qflow as a stand alone | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1456493 | CMRQflow_Standalone_3 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | NA | VbA | VbA | None
60 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1359246 | Common: Viewing or Analysis based on incorrect or wrong merging of split series | QFlow.Baseline.MV | 4.1 System Architecture | 1367609 | Qflow_HAZ_AppSpecific_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
61 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1350626 | qFlow: Results usability safety mitigations | QFlow.EARatio | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
62 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1350626 | qFlow: Results usability safety mitigations | QFlow.EARatio | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
63 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1361968 | qFlow: Showing overlay on underlay | QFlow.OverlayFusing | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
64 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1361968 | qFlow: Showing overlay on underlay | QFlow.OverlayFusing | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
65 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1350617 | qFlow : ROI validity usability safety mitigations | QFlow.ResultsMixUp | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
66 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1350616 | qFlow : Data validity usability safety mitigations | Qflow.InputDataQuality.Userawarness. | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
67 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1350617 | qFlow : ROI validity usability safety mitigations | Qflow.WrongContour.Userawarness. | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
68 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1350617 | qFlow : ROI validity usability safety mitigations | Qflow.WrongContour Propogation.Userawarness. | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
69 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1361969 | qFlow: ROI Lableing | Qflow.WrongLabel.Userawarness. | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
70 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1343190 | qFlow: Background correction usability safety mitigations | Qflow.BCContour Propogation.Userawarness. | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
71 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1350624 | qFlow: Graph usability safety mitigations | Qflow.DefineRange.Userawarness. | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
72 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | Qflow.ROIIdentify.TableGraph.Userawarness. | 4.2.1.1 Cardiac Suite (UI) | 1367609 | Qflow_HAZ_AppSpecific_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
73 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1350626 | qFlow: Results usability safety mitigations | Qflow.ROIIdentify.TableGraph.Userawarness. | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
74 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | Qflow.ROIIdentify.TableGraph.Userawarness. | 4.3.1 Cardiac.Common | 1367609 | Qflow_HAZ_AppSpecific_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
75 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | Qflow.ROIIdentify.TableGraph.Userawarness | 4.3.1 Cardiac.Common | 1367609 | Qflow_HAZ_AppSpecific_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
76 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | Qflow.BSAValues.Userawarness. | 4.3.1 Cardiac.Common | 1367609 | Qflow_HAZ_AppSpecific_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
77 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1300188 | Patient or Exam related parameters (user editable) - usability safety mitigations | Qflow.BSAValues.Userawarness. | 4.3.1 Cardiac.Common | 1367609 | Qflow_HAZ_AppSpecific_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
78 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1350626 | qFlow: Results usability safety mitigations | Qflow.K-Value.Userawarness | 4.3.5 QFlow Application | 1367609 | Qflow_HAZ_AppSpecific_16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | D000926820>3.3.5 | Pass | None | None
79 | 1354622 | MR Qflow | 1182555 | Qflow | Section 2.2.1 and Section 2.2.2 | 1300189 | Send results to finding dashboard - usability safety mitigations | Qflow.FindingsDashResultUpdate.Userawarness. | 4.3.1 Cardiac.Common | 1367609 | Qflow_HAZ_AppSpecific_17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
80 | 1179457 | Conversion to semantic model | 1436922 | Discrete Reporting | Section 2.2.1 and Section 2.2.2 | 1440093 | Discrete Reporting | QFlow.DiscreteReporting | 4.3.1 Cardiac.Common | 1367609 | Qflow_HAZ_AppSpecific_18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
81 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355342 | Common: Incorrect display of Pixel data due to 2D image manipulations | 4.3.2 | 4.1 System Architecture | 1367569 | Common_RMM_01 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
82 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355352 | Common: The user is scrolling images in a series using mouse click and drag. The user is scrolling in a speed that causes the system to skip slices. The user is unaware that Not all images have been shown | 4.3.12 | 4.1 System Architecture | 1367569 | Common_RMM_01 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
83 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355344 | Common: Wrong or missing contrast indication on displayed image | 4.3.4 | 4.1 System Architecture | 1367569 | Common_RMM_02 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
84 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355353 | Common: User loaded 2 series with different field of view. The “zoom link” is causing a part of the anatomy to be out of view The user think he is looking at the image in the original zoom | 4.3.14 | 4.1 System Architecture | 1367569 | Common_RMM_03 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
85 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355354 | Common: Showing wrong slice number under "phase" annotation in addition to the correct slice number under the "slice" annotation | 4.3.15 | 4.1 System Architecture | 1367569 | Common_RMM_04 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
86 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1359246 | Common: Viewing or Analysis based on incorrect or wrong merging of split series | 4.10. | 4.1 System Architecture | 1367569 | Common_RMM_5 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
87 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1356011 | Common: Geometrical measurement value or position is changed after activating a saved finding[4.4.7 ] | 4.4.7 | 4.1 System Architecture | 1366755 | MRC_RMF_1 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
88 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1356011 | Common: Geometrical measurement value or position is changed after activating a saved finding[4.4.7 ] | 4.4.7 | 4.1 System Architecture | 1366755 | MRC_RMF_1 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
89 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 4.4.8 | 4.1 System Architecture | 1366755 | MRC_RMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
90 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 4.4.8 | 4.1 System Architecture | 1366755 | MRC_RMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
91 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357606 | Common: Incorrect or incorrect presentation of scale, grid on images.[ 4.4.15] | 4.4.15 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_1 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
92 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357610 | Common: Textual annotation's position is changed due to image manipulations (e.g. zoom, pan, sweivel, rotate, flip)[ 4.4.20] | 4.4.20 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
93 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357620 | Common: Wrong interpretation of images when lossy compression is used 2. Wrong interpretation of images when lossy compression with resolution reduction is used[ 4.5.1] | 4.5.1 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_3 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
94 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357871 | Common: new images are saved to a wrong exam of the same patient | 4.7.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_1 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
95 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
96 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.2 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
97 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
98 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1, | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
99 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.2 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
100 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
101 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1, | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
102 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.2 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
103 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
104 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357897 | Common: Wrong correlation between images. Incorrect position of location represented in one or more of the viewports. | 4.11 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_5 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
105 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355340 | Common: Reversing of L/R, A/P, H/F. 2. Incorrect volume orientation | 4.3.1 | 4.1 System Architecture | 1367570 | Common_Orientation_RMM_01 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
106 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355340 | Common: Reversing of L/R, A/P, H/F. 2. Incorrect volume orientation | 4.3.1 | 4.1 System Architecture | 1367570 | Common_Orientation_RMM_02 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
107 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355340 | Common: Reversing of L/R, A/P, H/F. 2. Incorrect volume orientation | 4.3.1 | 4.1 System Architecture | 1367570 | Common_Orientation_RMM_03 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
108 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355340 | Common: Reversing of L/R, A/P, H/F. 2. Incorrect volume orientation | 4.3.1 | 4.1 System Architecture | 1367570 | Common_Orientation_RMM_04 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
109 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355340 | Common: Reversing of L/R, A/P, H/F. 2. Incorrect volume orientation | 4.3.1 | 4.1 System Architecture | 1367570 | Common_Orientation_RMM_05 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None
110 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355340 | Common: Reversing of L/R, A/P, H/F. 2. Incorrect volume orientation | 4.3.1 | 4.1 System Architecture | 1367570 | Common_Orientation_RMM_06 | Pass | NA | DHF380226>12.4 D000926820>3.3.5 | Pass | None | None

## TraceMatrix Spatial Enhancement
DESIGN INPUTS | [10 empty] | DESIGN VERIFICATION |  |  | DESIGN VALIDATION
SL No | URS Requirement ID | URS Feature (DHF380082,LH) | PRS Requirement ID | PRS Feature (DHF380211,Rev LJ ) | SDS ID (DHF380234 ,Rev LE) | DRS Requirement ID | DRS Feature (D000856146,Rev I) | RMM ID (D000879224 , Rev F) | DDS Id (D000993081,Rev D) | Verification Test Case (ID / Brief Description) (D000996623 ,Rev E) | Verification Test Step ID (D000996623 ,Rev E) | Verification Test Result (Pass/Fail) [Verification_Report: D000996308, Rev E] Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | Defect ID (for Failed steps from Verification) | Validation Test ID (Validation Plan DHF380226, Rev D, Validation Report D000926820 , Rev C) | Validation Test Result (Pass/Fail) [Validation Report D000926820, Rev C] | Defect ID (for Failed steps from Validation) | Remarks
1 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367012 | Data_Validation_01 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
2 | 1179490 | Suite Workflow | 1194223 | Semantic labeling: data identification | Section 2.2.1 and Section 2.2.2 | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.2 Functional Volumetric Application | 1367012 | Data_Validation_01 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
3 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367012 | Data_Validation_02 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
4 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367012 | Data_Validation_03 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
5 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367012 | Data_Validation_04 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
6 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367012 | Data_Validation_05 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
7 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367012 | Data_Validation_06 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
8 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296641 | Spatial - stages in application | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_01 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
9 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296451 | Contour creation and editing | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_02 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | None | None
10 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296451 | Contour creation and editing | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_03 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | None | None
11 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296451 | Contour creation and editing | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_04 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | None | None
12 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296451 | Contour creation and editing | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_05 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | None | None
13 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296452 | Segmenting scar tissue | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_06 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | None | None
14 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296452 | Segmenting scar tissue | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_07 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | None | None
15 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296452 | Segmenting scar tissue | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_08 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | None | None
16 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296452 | Segmenting scar tissue | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_09 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | None | None
17 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296452 | Segmenting scar tissue | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_10 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
18 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296452 | Segmenting scar tissue | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_11 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
19 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296453 | Spoke wheel | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_12 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
20 | 1353695 | MR Cardiac Functional LV & RV Analysis application | 1182556 | Functional LV/RV Analysis | Section 2.2.1 and Section 2.2.2 | 1285113 | Spokes wheel | NA | 4.3.2 Functional Volumetric Application | 1367035 | Worflow1_12 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
21 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296465 | Segmenting MVO tissue (no-reflow) | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_13 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
22 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296465 | Segmenting MVO tissue (no-reflow) | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_14 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
23 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296465 | Segmenting MVO tissue (no-reflow) | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_15 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
24 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296465 | Segmenting MVO tissue (no-reflow) | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_16 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
25 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296468 | Spatial - Result table | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_17 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
26 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296469 | Spatial - Polar maps | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_18 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
27 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238799 | Measurements: Measurement types | NA | 4.3.6 Cardiac (MR) Viewer | 1367035 | Worflow1_19 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | None | None
28 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1240463 | Attributes to be saved with the finding | NA | 4.3.2 Functional Volumetric Application | 1367035 | Worflow1_19 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | None | None
29 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1240463 | Attributes to be saved with the finding | NA | 4.3.2 Functional Volumetric Application | 1367035 | Worflow1_20 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
30 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1358644 | Application results as finding | NA | 4.3.2 Functional Volumetric Application | 1367035 | Worflow1_21 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
31 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1240468 | Manual capturing | NA | 4.3.2 Functional Volumetric Application | 1367035 | Worflow1_22 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
32 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1240465 | Objects that can be captured as a Finding | NA | 4.3.2 Functional Volumetric Application | 1367035 | Worflow1_23 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
33 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1318000 | Save findings | NA | 4.3.2 Functional Volumetric Application | 1367035 | Worflow1_24 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | VbA | VbA | None
34 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1440092 | Copy Table to clipboard | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_25 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | 1425115 | None
35 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_26 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
36 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367035 | Worflow1_27 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Vba | Vba | None
37 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367074 | Workflow2_01 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
38 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236005 | Viewing: Basic Viewing Capabilities | NA | 4.3.6 Cardiac (MR) Viewer | 1367074 | Workflow2_02 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | Vba | None
39 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236776 | Viewing: "View" menu | NA | 4.3.6 Cardiac (MR) Viewer | 1367074 | Workflow2_03 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | Vba | None
40 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1236776 | Viewing: "View" menu | NA | 4.3.6 Cardiac (MR) Viewer | 1367074 | Workflow2_04 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | Vba | None
41 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.6 Cardiac (MR) Viewer | 1367074 | Workflow2_05 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | Vba | None
42 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.6 Cardiac (MR) Viewer | 1367074 | Workflow2_06 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | Vba | None
43 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.6 Cardiac (MR) Viewer | 1367074 | Workflow2_07 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | Vba | None
44 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1239033 | Cine Widget | NA | 4.3.6 Cardiac (MR) Viewer | 1367074 | Workflow2_08 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | Vba | None
45 | 1179486 | Viewer application | 1182554 | Viewer | Section 2.2.1 and Section 2.2.2 | 1238741 | Viewing: Layout Management | NA | 4.3.6 Cardiac (MR) Viewer | 1367074 | Workflow2_09 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | Vba | None
46 | 1044734 | Filming | 529314 | ISPV7_PRS_ID_2_1: Film & Report | Section 2.2.1 and Section 2.2.3 | 1410882 | Save as option for images | NA | 4.3.2 Functional Volumetric Application | 1367074 | Workflow2_10 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | Vba | None
47 | 1044734 | Filming | 529314 | ISPV7_PRS_ID_2_1: Film & Report | Section 2.2.1 and Section 2.2.4 | 1413756 | Send to film (basic workflow) | NA | 4.3.2 Functional Volumetric Application | 1367074 | Workflow2_11 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | Vba | None
48 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1490331 | Import contours | NA | 4.3.4 Spatial Enhancement Application | 1367074 | Workflow2_12 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | 1425115 | None
49 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1490331 | Import contours | NA | 4.3.4 Spatial Enhancement Application | 1367074 | Workflow2_13 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.4 | Pass | 1425115 | None
50 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367074 | Workflow2_14 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | Vba | None
51 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367074 | Workflow2_15 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | Vba | None
52 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.4 Spatial Enhancement Application | 1367078 | Workflow3_01 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
53 | 1179490 | Suite Workflow | 1182550 | Multiple applications side by side | Section 2.2.1 and Section 2.2.2 | 1288854 | Displaying multiple applications side by side | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_02 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | Vba | None
54 | 1179490 | Suite Workflow | 1182550 | Multiple applications side by side | Section 2.2.1 and Section 2.2.2 | 1288854 | Displaying multiple applications side by side | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_03 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | Vba | None
55 | 1179488 | Integration to external applications | 1182552 | Launching external applications | Section 2.2.1 and Section 2.2.2 | 1233675 | Launching external applications | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_03 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | Vba | None
56 | 1179490 | Suite Workflow | 1182550 | Multiple applications side by side | Section 2.2.1 and Section 2.2.2 | 1288859 | Opening new applications | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_04 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | Vba | None
57 | 1179488 | Integration to external applications | 1182552 | Launching external applications | Section 2.2.1 and Section 2.2.2 | 1233675 | Launching external applications | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_04 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | Vba | None
58 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_05 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
59 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_05 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
60 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_06 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
61 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_06 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
62 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_07 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
63 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_07 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
64 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1233676 | Customized workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_08 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
65 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_08 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
66 | 1179490 | Suite Workflow | 1182550 | Multiple applications side by side | Section 2.2.1 and Section 2.2.2 | 1288859 | Opening new applications | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_09 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | 1425115 | None
67 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1233673 | Multiple Applications sides by side | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_09 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
68 | 1179490 | Suite Workflow | 1182550 | Multiple applications side by side | Section 2.2.1 and Section 2.2.2 | 1288854 | Displaying multiple applications side by side | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_09 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | 1425115 | None
69 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_10 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | Vba | None
70 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_11 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | Vba | None
71 | 1179490 | Suite Workflow | 1182551 | Customized workflow | Section 2.2.1 and Section 2.2.2 | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.2.1.2 Cardiac Suite (App) & Processing | 1367078 | Workflow3_12 | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226>12.4 | Pass | Vba | None
72 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1350658 | Spatial: Contour correctness usability safety mitigation | Spatial.IncorrectAuto Contour.Userawarness | 4.3.4 Spatial Enhancement Application | 1367093 | RMF_SPE_01 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
73 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1350658 | Spatial: Contour correctness usability safety mitigation | Spatial.IncorrectAuto Contour.Userawarness | 4.3.4 Spatial Enhancement Application | 1367093 | RMF_SPE_02 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
74 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1361474 | Spatial: Review and edit threshold based segmentation | Spatial.IncorrectEnhancement.Userawarness | 4.3.4 Spatial Enhancement Application | 1367093 | RMF_SPE_04 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
75 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1350666 | Spatial: Segment "no-reflow" (MVO) usability safety mitigations | Spatial.No-Reflow(MVO).Userawarness | 4.3.4 Spatial Enhancement Application | 1367093 | RMF_SPE_05 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
76 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1350661 | Spatial: Edit contours usability safety mitigation | Spatial.EditContour.IncorrectEnhancement.Userawarness | 4.3.4 Spatial Enhancement Application | 1367093 | RMF_SPE_06 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
77 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1350661 | Spatial: Edit contours usability safety mitigation | Spatial.EditContour.IncorrectEnhancement.Userawarness | 4.3.4 Spatial Enhancement Application | 1367093 | RMF_SPE_07 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
78 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | Spatial.EditContour.IncorrectEnhancement.Userawarness | 4.3.4 Spatial Enhancement Application | 1367093 | RMF_SPE_08 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
79 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | Spatial.IncorrectSpokeWheel.Userawarness | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_09 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
80 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | Spatial.IncorrectSpokeWheel.Userawarness | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_10 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
81 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | Spatial.IncorrectSpokeWheel.Userawarness | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_11 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
82 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | Spatial.IncorrectSpokeWheel.Userawarness | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_12 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
83 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | Spatial.IncorrectSpokeWheel.Userawarness | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_13 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
84 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300161 | Polar maps & spokes wheel - usability safety mitigations | Spatial.IncorrectSegment.SpokeWheel | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_14 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
85 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1350669 | Polar maps & spokes wheel - usability safety mitigations | Spatial.IncorrectSegment.SpokeWheel | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_15 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
86 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | Spatial.IncorrectMeasurement LinkTabel.Userawarness. | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_15 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
87 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | Spatial.IncorrectMeasurement LinkTabel.Userawarness. | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_15 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
88 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1350669 | Spatial: Results review usability safety mitigations | Spatial.IncorrectMeasurement LinkTabel.Userawarness. | 4.3.4 Spatial Enhancement Application | 1367093 | RMF_SPE_16 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
89 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | Spatial.IncorrectMeasurement LinkTabel.Userawarness. | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_16 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
90 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | Spatial.IncorrectMeasurement LinkTabel.Userawarness. | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_16 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
91 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1350669 | Spatial: Results review usability safety mitigations | Spatial.IncorrectMeasurement LinkTabel.Userawarness. | 4.3.4 Spatial Enhancement Application | 1367093 | RMF_SPE_17 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
92 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300181 | Review the measurements in the tables - usability safety mitigations | Spatial.IncorrectMeasurement LinkTabel.Userawarness. | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_17 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
93 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | Spatial.IncorrectMeasurement LinkTabel.Userawarness. | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_17 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | None
94 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.1 Cardiac.Common | 1367012 | Data_Validation_01 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | NA
95 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.1 Cardiac.Common | 1367012 | Data_Validation_02 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | NA
96 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.1 Cardiac.Common | 1367012 | Data_Validation_03 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | NA
97 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.1 Cardiac.Common | 1367012 | Data_Validation_04 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | NA
98 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.1 Cardiac.Common | 1367012 | Data_Validation_05 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | NA
99 | 1353733 | MR Cardiac Spatial Enhancement Analysis application | 1182559 | Spatial Enhancement | Section 2.2.1 and Section 2.2.2 | 1296450 | Spatial - Supported data | NA | 4.3.1 Cardiac.Common | 1367012 | Data_Validation_06 | Pass | NA | DHF380226 >12.4 D000926820>3.3.6 | Pass | 1425115 | NA
100 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1300189 | Send results to finding dashboard - usability safety mitigations | Spatial..FindingsDashResultUpdate.Userawarness. | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_18 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
101 | 1179457 | Conversion to semantic model | 1436922 | Discrete Reporting | Section 2.2.1 and Section 2.2.2 | 1440093 | Discrete Reporting | Spatial.DiscreteReporting | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_19 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
102 | 1179457 | Conversion to semantic model | 1436922 | Discrete Reporting | Section 2.2.1 and Section 2.2.2 | 1440093 | Discrete Reporting | Spatial.DiscreteReporting | 4.3.1 Cardiac.Common | 1367093 | RMF_SPE_20 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
103 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355342 | Common: Incorrect display of Pixel data due to 2D image manipulations | 4.3.2 | 4.1 System Architecture | 1367569 | Common_RMM_1 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
104 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355352 | Common: The user is scrolling images in a series using mouse click and drag. The user is scrolling in a speed that causes the system to skip slices. The user is unaware that Not all images have been shown | 4.3.12 | 4.1 System Architecture | 1367569 | Common_RMM_1 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
105 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355344 | Common: Wrong or missing contrast indication on displayed image | 4.3.4 | 4.1 System Architecture | 1367569 | Common_RMM_2 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
106 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355353 | Common: User loaded 2 series with different field of view. The “zoom link” is causing a part of the anatomy to be out of view The user think he is looking at the image in the original zoom | 4.3.14 | 4.1 System Architecture | 1367569 | Common_RMM_3 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
107 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355354 | Common: Showing wrong slice number under "phase" annotation in addition to the correct slice number under the "slice" annotation | 4.3.15 | 4.1 System Architecture | 1367569 | Common_RMM_4 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
108 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1359246 | Common: Viewing or Analysis based on incorrect or wrong merging of split series | 4.10. | 4.1 System Architecture | 1367569 | Common_RMM_5 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
109 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1356011 | Common: Geometrical measurement value or position is changed after activating a saved finding[4.4.7 ] | 4.4.7 | 4.1 System Architecture | 1366755 | MRC_RMF_1 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
110 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 4.4.8 | 4.1 System Architecture | 1366755 | MRC_RMF_2 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
111 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357606 | Common: Incorrect or incorrect presentation of scale, grid on images.[ 4.4.15] | 4.4.15 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_1 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
112 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357610 | Common: Textual annotation's position is changed due to image manipulations (e.g. zoom, pan, sweivel, rotate, flip)[ 4.4.20] | 4.4.20 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_2 | NA | NA | DHF380226>12.4 | Vba | Vba | This Test step was designed for QFlow analysis application, hence not applicable for Spatial Enhancement
113 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357620 | Common: Wrong interpretation of images when lossy compression is used 2. Wrong interpretation of images when lossy compression with resolution reduction is used[ 4.5.1] | 4.5.1 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_3 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
114 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357871 | Common: new images are saved to a wrong exam of the same patient | 4.7.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_1 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
115 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
116 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
117 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
118 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | Pass | NA | DHF380226>12.4 | Vba | Vba | None
119 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357897 | Common: Wrong correlation between images. Incorrect position of location represented in one or more of the viewports. | 4.11 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_5 | Pass | NA | DHF380226>12.4 | Vba | Vba | None

## Cardiac Mapping_Traceability
Design Inputs | [8 empty] | Design Verification |  |  |  |  | Design Validation
URS Requirement ID | URS Feature (DHF380082,LH) | PRS Requirement ID | PRS Feature (DHF380211,Rev LJ ) | SDS ID (DHF380234 ,Rev LE) | DRS Requirement ID | DRS Feature (D000856146,Rev I) | RMM ID (D000879224 , Rev F) | DDS Id (D000993081,Rev D) | Test Case ID Ref: | Test Case Name | Test Step details | Verification Test Result (Pass/Fail) [Verification_Report: D000996308, Rev E] Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | Defect ID (for Failed steps from Verification) | Validation Test ID (Validation Plan DHF380226, Rev D, Validation Report D000926820 , Rev C) | Validation Test Result (Pass/Fail) [Validation Report D000926820, Rev C] | Defect ID (for Failed steps from Validation) | Remarks
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328216 | Datatype | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240929 | Viewer: Factory hanging protocol (based on the loaded data) | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328579 | Application launch | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328582 | Map creation and motion correction stage | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395896 | Hanging protocol (Displayed viewports) | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395900 | Generate map and error map | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396974 | Motion correction and excluding data points | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396974 | Motion correction and excluding data points | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395908 | Quality control tools | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395908 | Quality control tools | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395896 | Hanging protocol (Displayed viewports) | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396976 | Select map for assessment | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328597 | Draw and edit LV Endo and Epi contours contours | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328597 | Draw and edit LV Endo and Epi contours contours | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346487 | Colormap default range | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328601 | Draw and edit blood pool for ECV calculation | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346524 | ECV calculation | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396989 | Set Hematocrit for ECV calculation | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346524 | ECV calculation | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 20 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 20 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346439 | Spokes wheel in Mapping | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 21 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 21 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440092 | Copy Table to clipboard | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 22 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328597 | Draw and edit LV Endo and Epi contours contours | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 23 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328597 | Draw and edit LV Endo and Epi contours contours | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 24 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328597 | Draw and edit LV Endo and Epi contours contours | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 25 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328597 | Draw and edit LV Endo and Epi contours contours | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 26 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328597 | Draw and edit LV Endo and Epi contours contours | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 27 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328597 | Draw and edit LV Endo and Epi contours contours | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 28 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328597 | Draw and edit LV Endo and Epi contours contours | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 29 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 30 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346487 | Colormap default range | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 30 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1358644 | Application results as finding | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 31 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240468 | Manual capturing | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 32 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240465 | Objects that can be captured as a Finding | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 33 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240465 | Objects that can be captured as a Finding | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 34 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 35 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 36 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 37 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420421 | Cardiac Mapping-T1 Native & Enhanced Series workflow | Step 38 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328216 | Datatype | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240929 | Viewer: Factory hanging protocol (based on the loaded data) | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240929 | Viewer: Factory hanging protocol (based on the loaded data) | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328579 | Application launch | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328582 | Map creation and motion correction stage | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395900 | Generate map and error map | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328582 | Map creation and motion correction stage | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396976 | Select map for assessment | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346496 | T1 table result | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240468 | Manual capturing | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1358644 | Application results as finding | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1413756 | Send to film (basic workflow) | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420422 | Cardiac Mapping-T1 Native MOLLI Workflow | Step 16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328216 | Datatype | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328579 | Application launch | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395900 | Generate map and error map | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328582 | Map creation and motion correction stage | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395896 | Hanging protocol (Displayed viewports) | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328582 | Map creation and motion correction stage | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395908 | Quality control tools | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396976 | Select map for assessment | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346442 | Colormaps | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications |  |  | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346489 | Result types | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346495 | T2 Table results | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346489 | Result types | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328597 | Draw and edit LV Endo and Epi contours contours | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346439 | Spokes wheel in Mapping | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240468 | Manual capturing | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1318000 | Save findings | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420423 | Cardiac Mapping-T1 Native ShMOLLI workflow | Step 17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328216 | Datatype | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328579 | Application launch | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328581 | Application stages | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328582 | Map creation and motion correction stage | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395896 | Hanging protocol (Displayed viewports) | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395900 | Generate map and error map | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395912 | Fitted Curve graph | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395908 | Quality control tools | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396976 | Select map for assessment | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346442 | Colormaps | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328133 | General | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1358644 | Application results as finding | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1318000 | Save findings | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420424 | Cardiac Mapping-T1 Native SASHA workflow | Step 16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328216 | Datatype | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328216 | Datatype | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328579 | Application launch | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395900 | Generate map and error map | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396974 | Motion correction and excluding data points | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396976 | Select map for assessment | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346442 | Colormaps | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328599 | Regional (Local) segmentation | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346509 | Graph | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346509 | Graph | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240468 | Manual capturing | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1358644 | Application results as finding | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440093 | Discrete Reporting: | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | VbA | VbA | None
1179490 | Suite Workflow | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440093 | Discrete Reporting: | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | VbA | VbA | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440093 | Discrete Reporting: | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1179490 | Suite Workflow | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440093 | Discrete Reporting: | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1318000 | Save findings | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1318000 | Save findings | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346439 | Spokes wheel in Mapping | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 20 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 21 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1358644 | Application results as finding | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 22 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240468 | Manual capturing | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 23 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1318000 | Save findings | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 24 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 25 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 26 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420425 | Cardiac Mapping-T2 Series workflow | Step 27 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328216 | Datatype | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328216 | Datatype | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237750 | Semantic labeling: Semantic labeling for all modalities | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328579 | Application launch | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395900 | Generate map and error map | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328582 | Map creation and motion correction stage | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395912 | Fitted Curve graph | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396974 | Motion correction and excluding data points | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396976 | Select map for assessment | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346442 | Colormaps | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Draw ROI on blood/muscle/liver | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328599 | Regional (Local) segmentation | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328602 | Draw ROI on blood/muscle/liver | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346495 | T2 Table results | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346509 | Graph | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346509 | Graph | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1346495 | T2 Table results | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328596 | Segmentation & Result stage default hanging protocol and image position | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440092 | Copy Table to clipboard | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications |  |  | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1358644 | Application results as finding | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1240468 | Manual capturing | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 20 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1318000 | Save findings | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 21 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 22 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 23 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | 4.3.7 Mapping Application | 1420426 | Cardiac Mapping-T2* Series workflow | Step 24 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328216 | Datatype | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328579 | Application launch | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328581 | Application stages | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328581 | Application stages | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412554 | Import contours from functional LV | NA | 4.3.7 Mapping Application | 1474095 | Cardiac Mapping-Import Contours | Step 15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1465661 | synthetic ECV calculation from synthetic HCT. | NA | 4.3.7 Mapping Application | 1490253 | Synthetic ECV Calculation | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396974 | Motion correction and excluding data points | NA | 4.3.7 Mapping Application | 1490253 | Synthetic ECV Calculation | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396976 | Select map for assessment | NA | 4.3.7 Mapping Application | 1490253 | Synthetic ECV Calculation | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328581 | Application stages | NA | 4.3.7 Mapping Application | 1490253 | Synthetic ECV Calculation | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1396974 | Motion correction and excluding data points | NA | 4.3.7 Mapping Application | 1490253 | Synthetic ECV Calculation | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328601 | Draw and edit blood pool for ECV calculation | NA | 4.3.7 Mapping Application | 1490253 | Synthetic ECV Calculation | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1465661 | synthetic ECV calculation from synthetic HCT. | NA | 4.3.7 Mapping Application | 1490253 | Synthetic ECV Calculation | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1465661 | synthetic ECV calculation from synthetic HCT. | NA | 4.3.7 Mapping Application | 1490253 | Synthetic ECV Calculation | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1465661 | synthetic ECV calculation from synthetic HCT. | NA | 4.3.7 Mapping Application | 1490253 | Synthetic ECV Calculation | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 1179490 | MR Cardiac Quantitative Mapping application Suite Workflow | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.3.7 Mapping Application | 1490254 | Cardiac Mapping: Customized Workflow | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 1179490 | MR Cardiac Quantitative Mapping application Suite Workflow | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357906 | Workflow : Create new workflow | NA | 4.3.7 Mapping Application | 1490254 | Cardiac Mapping: Customized Workflow | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 1179490 | MR Cardiac Quantitative Mapping application Suite Workflow | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357906 | Workflow : Create new workflow | NA | 4.3.7 Mapping Application | 1490254 | Cardiac Mapping: Customized Workflow | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 1179490 | MR Cardiac Quantitative Mapping application Suite Workflow | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357906 | Workflow : Create new workflow | NA | 4.3.7 Mapping Application | 1490254 | Cardiac Mapping: Customized Workflow | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 1179490 | MR Cardiac Quantitative Mapping application Suite Workflow | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357906 | Workflow : Create new workflow | NA | 4.3.7 Mapping Application | 1490254 | Cardiac Mapping: Customized Workflow | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 1179490 | MR Cardiac Quantitative Mapping application Suite Workflow | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357906 | Workflow : Create new workflow | NA | 4.3.7 Mapping Application | 1490254 | Cardiac Mapping: Customized Workflow | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 1179490 | MR Cardiac Quantitative Mapping application Suite Workflow | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.3.7 Mapping Application | 1490254 | Cardiac Mapping: Customized Workflow | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328216 | Datatype | NA | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328579 | Application launch | NA | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237750 | Semantic labeling: Semantic labeling for all modalities | Mapping.MotionCorrection.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328579 | Application launch | Mapping.MotionCorrection.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411055 | Apply motion correction | Mapping.MotionCorrection.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411046 | Verify map quality (in Re-calculation stage) | Mapping.MotionCorrection.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411046 | Verify map quality (in Re-calculation stage) | Mapping.MotionCorrection.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411055 | Apply motion correction | Mapping.MotionCorrection.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411055 | Apply motion correction | Mapping.MotionCorrection.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1300186 | Review graph - usability safety mitigations | Mapping.MotionCorrection.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1300187 | Long text indications - usability safety mitigations | Mapping.MotionCorrection.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411055 | Apply motion correction | Mapping.MotionCorrection.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411060 | Contours accuracy | Mapping.InaccurateContours.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411062 | Color map | Mapping.InaccurateContours.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411060 | Contours accuracy | Mapping.InaccurateContours.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411060 | Contours accuracy | Mapping.InaccurateContours.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411060 | Contours accuracy | Mapping.InaccurateContours.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1411062 | Color map | Mapping.T1ColorMapping.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1328601 | Draw and edit blood pool for ECV calculation | Mapping.TabelMeasurement.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1300187 | Long text indications - usability safety mitigations | Mapping.TabelMeasurement.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1300161 | Polar maps & spokes wheel - usability safety mitigations | Mapping.PolarMaps.Userawarness | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1300181 | Review the measurements in the tables - usability safety mitigations | Mapping.TabelMeasurement.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A |  |  | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1300187 | Long text indications - usability safety mitigations | Mapping.TabelMeasurement.Userawarness. | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1300161 | Polar maps & spokes wheel - usability safety mitigations | Mapping.PolarMaps.Userawarness | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1300161 | Polar maps & spokes wheel - usability safety mitigations | Mapping.PolarMaps.Userawarness | 4.3.7 Mapping Application | 1420622 | Cardiac Mapping-RMF Test Case | Step 20 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1300189 | Send results to finding dashboard - usability safety mitigations | Mapping.FindingsDashResultUpdate.Userawarness. | 4.3.7 Mapping Application | 1366755 | TC_Common Hazard_Send results to finding dashboard | Step 1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1300189 | Send results to finding dashboard - usability safety mitigations | Mapping.FindingsDashResultUpdate.Userawarness. | 4.3.7 Mapping Application | 1366755 | TC_Common Hazard_Send results to finding dashboard | Step 2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357891 | Common: The saved finding doesn't appear where user expects it [ 4.8.7] | 4.8.7 Mapping.FindingsDashResultUpdate.Userawarness. | 4.3.7 Mapping Application | 1366755 | TC_Common Hazard_Send results to finding dashboard | Step 2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357613 | Common: Finding preview image is wrong or inaccurate[ 4.4.26] | 4.4.26 Mapping.FindingsDashResultUpdate.Userawarness. | 4.3.7 Mapping Application | 1366755 | TC_Common Hazard_Send results to finding dashboard | Step 2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357606 | Common: Incorrect or incorrect presentation of scale, grid on images.[ 4.4.15] | 4.4.15 | 4.3.7 Mapping Application | 1366785 | TC_Measurements common_incorrect presentation of scale, grid on images | Step 1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357610 | Common: Textual annotation's position is changed due to image manipulations (e.g. zoom, pan, sweivel, rotate, flip)[ 4.4.20] | 4.4.20 | 4.3.7 Mapping Application | 1366785 | TC_Measurements common_incorrect presentation of scale, grid on images | Step 2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357610 | Common: Textual annotation's position is changed due to image manipulations (e.g. zoom, pan, sweivel, rotate, flip)[ 4.4.20] | 4.4.20 | 4.3.7 Mapping Application | 1366785 | TC_Measurements common_incorrect presentation of scale, grid on images | Step 3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357871 | Common: new images are saved to a wrong exam of the same patient[ 4.7.1] | 4.7.1 | 4.3.7 Mapping Application | 1367400 | Common Hazard TC_Cine Movie and New images saved | Step 1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357871 | Common: new images are saved to a wrong exam of the same patient[ 4.7.1] | 4.7.1 | 4.3.7 Mapping Application | 1367400 | Common Hazard TC_Cine Movie and New images saved | Step 2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate..[ 4.8.1, 4.8.2, 4.8.3] | 4.8.1 | 4.3.7 Mapping Application | 1367400 | Common Hazard TC_Cine Movie and New images saved | Step 3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate..[ 4.8.1, 4.8.2, 4.8.3] | 4.8.1 | 4.3.7 Mapping Application | 1367400 | Common Hazard TC_Cine Movie and New images saved | Step 4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate..[ 4.8.1, 4.8.2, 4.8.3] | 4.8.1 | 4.3.7 Mapping Application | 1367400 | Common Hazard TC_Cine Movie and New images saved | Step 5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1356015 | Common: Line measurements' values changed due to Zoom[4.4.9.2 ] | 4.4.9.2 | 4.3.7 Mapping Application | 1490278 | Common Hazard : Line measurements & Geometrical values changed due to Zoom | Step 1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1357584 | Common: Pixel : Geometrical measurements' values changed with more than allowed error due to image manipulations (zoom, pan, rotate, mirror, invert, sharpness)[4.4.9.5] | 4.4.9.5 | 4.3.7 Mapping Application | 1490278 | Common Hazard : Line measurements & Geometrical values changed due to Zoom | Step 2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412555 | Mapping normal values | NA | 4.3.7 Mapping Application | 1420579 | CardiacMapping-T1 Native & Enhanced SanityTestCase | Step 13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1272000 | [CLM] MR cardiac suite | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412555 | Mapping normal values | NA | 4.3.7 Mapping Application | 1420579 | CardiacMapping-T1 Native & Enhanced SanityTestCase | Step 13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412555 | Mapping normal values | NA | 4.3.7 Mapping Application | 1420579 | CardiacMapping-T1 Native & Enhanced SanityTestCase | Step 14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1272000 | [CLM] MR cardiac suite | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412555 | Mapping normal values | NA | 4.3.7 Mapping Application | 1420579 | CardiacMapping-T1 Native & Enhanced SanityTestCase | Step 14 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412555 | Mapping normal values | NA | 4.3.7 Mapping Application | 1420579 | CardiacMapping-T1 Native & Enhanced SanityTestCase | Step 15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1272000 | [CLM] MR cardiac suite | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412555 | Mapping normal values | NA | 4.3.7 Mapping Application | 1420579 | CardiacMapping-T1 Native & Enhanced SanityTestCase | Step 15 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412555 | Mapping normal values | NA | 4.3.7 Mapping Application | 1420579 | CardiacMapping-T1 Native & Enhanced SanityTestCase | Step 16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1272000 | [CLM] MR cardiac suite | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412555 | Mapping normal values | NA | 4.3.7 Mapping Application | 1420579 | CardiacMapping-T1 Native & Enhanced SanityTestCase | Step 16 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1353684 | MR Cardiac Quantitative Mapping application | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412555 | Mapping normal values | NA | 4.3.7 Mapping Application | 1420579 | CardiacMapping-T1 Native & Enhanced SanityTestCase | Step 17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None
1272000 | [CLM] MR cardiac suite | 1182560 | Mapping | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412555 | Mapping normal values | NA | 4.3.7 Mapping Application | 1420579 | CardiacMapping-T1 Native & Enhanced SanityTestCase | Step 17 | PASS | NA | DHF380226 >12.4 D000926820>3.3.9 | Pass | None | None

## TemporalEnhancement_Tracebility
DESIGN INPUTS | [10 empty] | DESIGN VERIFICATION |  |  | DESIGN VALIDATION
SL No | URS Requirement ID | URS Feature (DHF380082,LH) | PRS Requirement ID | PRS Feature (DHF380211,Rev LJ ) | SDS ID (DHF380234 ,Rev LE) | DRS Requirement ID | DRS Feature (D000856146,Rev I) | RMM ID (D000879224 , Rev F) | DDS Id (D000993081,Rev D) | Verification Test Case (ID / Brief Description) (D001104277 ,Rev F) | Verification Test Step ID (D001104277 ,Rev F) | Verification Test Result (PASS/Fail) [Verification_Report: D000996308, Rev E] Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | Defect ID (for Failed steps from Verification) | Validation Test ID (Validation Plan DHF380226, Rev D, Validation Report D000926820 , Rev C) | Validation Test Result (Pass/Fail) [Validation Report D000926820, Rev C] | Defect ID (for Failed steps from Validation) | Remarks
1 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410561 | Supported datatypes and sequences | NA | 4.3.6. Temporal Enhancement Application | 5.1: Data Loading (1423675) | Launch _2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
2 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410561 | Supported datatypes and sequences | NA | 4.3.6. Temporal Enhancement Application | 5.1: Data Loading (1423675) | Launch _3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
3 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410561 | Supported datatypes and sequences | NA | 4.3.6. Temporal Enhancement Application | 5.1: Data Loading (1423675) | Launch _5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
4 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410561 | Supported datatypes and sequences | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
5 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410561 | Supported datatypes and sequences | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
6 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.1: Data Loading (1423675) | Launch _4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
7 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.1: Data Loading (1423675) | Launch_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
8 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.1: Data Loading (1423675) | Launch _6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
9 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.1: Data Loading (1423675) | Launch _7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
10 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.1: Data Loading (1423675) | Launch _8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
11 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.1: Data Loading (1423675) | Launch _9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
12 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.1: Data Loading (1423675) | Launch _10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
13 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
14 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
15 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
16 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
17 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | Loading the application | NA | 4.3.6. Temporal Enhancement Application | 5.1.7:Sanity workflow(1424163) | App_RMF_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
18 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410563 | Hanging Protocols | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
19 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410563 | Hanging Protocols | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
20 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410563 | Hanging Protocols | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
21 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410563 | Hanging Protocols | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
22 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410563 | Hanging Protocols | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
23 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410564 | Contours | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
24 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410564 | Contours | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
25 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410564 | Contours | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
26 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410564 | Contours | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
27 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410564 | Contours | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
28 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410564 | Contours | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | 1456029 | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
29 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410564 | Contours | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
30 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410564 | Contours | NA | 4.3.6. Temporal Enhancement Application | 5.3: Application workflow with hanging protocol and contours. (1423681) | HP_Contour_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
31 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410568 | Dynamic range selection | Temp.WrongBloodPoolSegmentation.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
32 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410568 | Dynamic range selection | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
33 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410568 | Dynamic range selection | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
34 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410568 | Dynamic range selection | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
35 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410568 | Dynamic range selection | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
36 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410568 | Dynamic range selection | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
37 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410571 | Registration | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
38 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410571 | Registration | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
39 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410571 | Registration | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
40 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410571 | Registration | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
41 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410571 | Registration | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
42 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410571 | Registration | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
43 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410571 | Registration | NA | 4.3.6. Temporal Enhancement Application | 5.2: Application workflow with Dynamic range selection and Dynamic registration (1423676) | Dy-Reg-12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
44 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
45 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
46 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
47 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
48 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
49 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
50 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
51 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
52 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
53 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_11 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
54 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_12 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
55 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410573 | Graph | NA | 4.3.6. Temporal Enhancement Application | 5.6: Application workflow with Graph. (1424160) | Graph_13 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
56 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410574 | Measurements | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_9 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | Testcase and DRS needs to be updated as calculation needs to be removed from requirement and testcase.
57 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410574 | Measurements | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
58 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410575 | Tables | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
59 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410575 | Tables | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
60 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410575 | Tables | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
61 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410575 | Tables | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
62 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410575 | Tables | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
63 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410575 | Tables | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_10 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
64 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410575 | Tables | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
65 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410575 | Tables | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
66 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1440092 | Copy Table to clipboard | NA | 4.3.6. Temporal Enhancement Application | 5.4: Application workflow with Table (1424143) | Table_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
67 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
68 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
69 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
70 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
71 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
72 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
73 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
74 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
75 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
76 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
77 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
78 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410576 | Polar plots | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
79 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
80 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
81 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
82 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
83 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
84 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
85 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
86 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
87 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
88 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
89 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410579 | Perfusion fusion image | NA | 4.3.6. Temporal Enhancement Application | 5.5: Application workflow with Polar Plot and Perfusion Fusion Image.(1424155) | Polar_Per_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
90 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410562 | 5.2.11.6 Temporal Enhancment | NA | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_1 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
91 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410566 | 5.2.11.6 Temporal Enhancment | Temp.WrongBloodPoolSegmentation.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_2 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
92 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410566 | 5.2.11.6 Temporal Enhancment | Temp.WrongBloodPoolSegmentation.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_3 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
93 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410566 | 5.2.11.6 Temporal Enhancment | Temp.WrongBloodPoolSegmentation.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.1.2 Application workflow with Dynamic range selection and Dynamic registration | Dy-Reg-2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
94 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410569 | 5.2.11.6 Temporal Enhancment | Temp.WrongDefaultDynamicRange.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_4 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
95 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410570 | 5.2.11.6 Temporal Enhancment | Temp.NoMotionCorrection.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_5 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
96 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410570 | 5.2.11.6 Temporal Enhancment | Temp.NoMotionCorrection.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_6 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
97 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410570 | 5.2.11.6 Temporal Enhancment | Temp.NoMotionCorrection.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_7 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
98 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410570 | 5.2.11.6 Temporal Enhancment | Temp.NoMotionCorrection.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_8 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
99 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410572 | 5.2.11.6 Temporal Enhancment | Temporal.Graph.Userawarness | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_9 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
100 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410572 | 5.2.11.6 Temporal Enhancment | Temporal.Graph.Userawarness | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_10 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
101 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410572 | 5.2.11.6 Temporal Enhancment | Temporal.Graph.Userawarness | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_11 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
102 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410572 | 5.2.11.6 Temporal Enhancment | Temporal.Graph.Userawarness | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_12 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
103 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410577 | 5.2.11.6 Temporal Enhancment | Temporal.TabelMeasurement.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_13 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
104 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1300181 | 5.3.9.1.2 Review the measurements in the tables - usability safety mitigations | Temporal.TabelMeasurement.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
105 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410577 | 5.2.11.6 Temporal Enhancment | Temporal.TabelMeasurement.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_14 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
106 |  |  |  |  | Section 2.2.1 and Section 2.2.2 | 1457111 |  |  | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A |  |  | Pass
107 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410577 | 5.2.11.6 Temporal Enhancment | Temporal.TabelMeasurement.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
108 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410577 | 5.2.11.6 Temporal Enhancment | Temporal.TabelMeasurement.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
109 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410577 | 5.2.11.6 Temporal Enhancment | Temporal.TabelMeasurement.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
110 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410580 | 5.2.11.6 Temporal Enhancment | Temporal.TabelMeasurement.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
111 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410577 | 5.2.11.6 Temporal Enhancment | Temporal.TabelMeasurement.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
112 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410580 | 5.2.11.6 Temporal Enhancment | Temporal.TabelMeasurement.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
113 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1300186 | 5.2.11.6 Temporal Enhancment | Temporal.Graph.Userawarness | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
114 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1300189 | 5.2.11.6 Temporal Enhancment | Temporal.FindingsDashResultUpdate.Userawarness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_20 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
115 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1440093 | 5.2.11.6 Temporal Enhancment | Temporal.DiscreteReporting | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_21 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
116 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1440093 | 5.2.11.6 Temporal Enhancment | Temporal.DiscreteReporting | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_22 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
117 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1438347 | 5.3.7.9 Result Summary card | NA | 4.3.6. Temporal Enhancement Application | 5.1.5: Application workflow with Polar Plot and Perfusion Fusion Image. | Polar_Per_15 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | Finding dashboard functionality is supported in RC release of ISP13 for Temporal Enhancement application.
118 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1438347 | 5.3.7.9 Result Summary card | NA | 4.3.6. Temporal Enhancement Application | 5.1.5: Application workflow with Polar Plot and Perfusion Fusion Image. | Polar_Per_16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | Finding dashboard functionality is supported in RC release of ISP13 for Temporal Enhancement application.
119 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1412625 | 5.3.10.2 Reliability | NA | 4.3.6. Temporal Enhancement Application | 5.1.6: Non functional workflow | Non-funct_1 | Fail | 1497348 | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
120 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1412625 | 5.3.10.2 Reliability | NA | 4.3.6. Temporal Enhancement Application | 5.1.6: Non functional workflow | Non-funct_2 | NA | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | This test step cannot be executed as it is dependent on test step Non-funct_1 which is failing.
121 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1412625 | 5.3.10.2 Reliability | NA | 4.3.6. Temporal Enhancement Application | 5.1.6: Non functional workflow | Non-funct_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
122 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1412625 | 5.3.10.2 Reliability | NA | 4.3.6. Temporal Enhancement Application | 5.1.6: Non functional workflow | Non-funct_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
123 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | NA | 4.3.6. Temporal Enhancement Application | 5.1.6: Non functional workflow | Non-funct_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
124 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1300187 | Long text indications - usability safety mitigations | NA | 4.3.6. Temporal Enhancement Application | 5.1.6: Non functional workflow | Non-funct_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
125 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410580 | 5.3.9.6.7 Perfusion fusion images | Temporal.Fusion.Userawareness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | NA
126 | 1353689 | MR Cardiac Temporal Enhancement | 1182558 | Temporal Enhancement | Section 2.2.1 and Section 2.2.2 | 1410580 | 5.3.9.6.7 Perfusion fusion images | Temporal.Fusion.Userawareness. | 4.3.6. Temporal Enhancement Application | 5.2.6:Application Hazard (1424524) | App_RMF_18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | NA
127 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355342 | Common: Incorrect display of Pixel data due to 2D image manipulations | 4.3.2 | 4.1 System Architecture | 1367569 | Common_RMM_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
128 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355352 | Common: The user is scrolling images in a series using mouse click and drag. The user is scrolling in a speed that causes the system to skip slices. The user is unaware that Not all images have been shown | 4.3.12 | 4.1 System Architecture | 1367569 | Common_RMM_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
129 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355344 | Common: Wrong or missing contrast indication on displayed image | 4.3.4 | 4.1 System Architecture | 1367569 | Common_RMM_2 | NA | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | Contrast agent is not supported in Temporal Enhancement application.
130 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355353 | Common: User loaded 2 series with different field of view. The “zoom link” is causing a part of the anatomy to be out of view The user think he is looking at the image in the original zoom | 4.3.14 | 4.1 System Architecture | 1367569 | Common_RMM_3 | NA | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | This test step is not applicable to Temporal Enhancement application.
131 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1355354 | Common: Showing wrong slice number under "phase" annotation in addition to the correct slice number under the "slice" annotation | 4.3.15 | 4.1 System Architecture | 1367569 | Common_RMM_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
132 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1359246 | Common: Viewing or Analysis based on incorrect or wrong merging of split series | 4.10. | 4.1 System Architecture | 1367569 | Common_RMM_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
133 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1356011 | Common: Geometrical measurement value or position is changed after activating a saved finding[4.4.7 ] | 4.4.7 | 4.1 System Architecture | 1366755 | MRC_RMF_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
134 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1356011 | Common: Geometrical measurement value or position is changed after activating a saved finding[4.4.7 ] | 4.4.7 | 4.1 System Architecture | 1366755 | MRC_RMF_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
135 | 1179490 | Suite Workflow | 1182563 | Displaying findings from all applications | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 4.4.8 | 4.1 System Architecture | 1366755 | MRC_RMF_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
136 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1356013 | Common: Wrong measurements displayed during editing of annotations[4.4.8 ] | 4.4.8 | 4.1 System Architecture | 1366755 | MRC_RMF_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
137 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357606 | Common: Incorrect or incorrect presentation of scale, grid on images.[ 4.4.15] | 4.4.15 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
138 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357610 | Common: Textual annotation's position is changed due to image manipulations (e.g. zoom, pan, sweivel, rotate, flip)[ 4.4.20] | 4.4.20 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_2 | NA | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | This test step is not applicable in Temporal Enhancement application
139 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357620 | Common: Wrong interpretation of images when lossy compression is used 2. Wrong interpretation of images when lossy compression with resolution reduction is used[ 4.5.1] | 4.5.1 | 4.1 System Architecture | 1366785 | MRC_CommonRMF_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
140 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357871 | Common: new images are saved to a wrong exam of the same patient | 4.7.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
141 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
142 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_3 | PASS |  | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
143 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.3 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | PASS |  | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
144 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357886 | Common: Cine movies are skipping frames to maintain the acquisition or clinical frame rate. | 4.8.1 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
145 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357897 | Common: Wrong correlation between images. Incorrect position of location represented in one or more of the viewports. | 4.11 | 4.1 System Architecture | 1367400 | Generic_AllApps_RMF_5 | NA | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | Relate functionality is not supported in Temporal Enhancement application.
146 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1490330 | Import Contours | NA | 4.3.6. Temporal Enhancement Application | 1423681 | HP_Contour_16 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
147 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1490330 | Import Contours | NA | 4.3.6. Temporal Enhancement Application | 1423681 | HP_Contour_17 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
148 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1490330 | Import Contours | NA | 4.3.6. Temporal Enhancement Application | 1423681 | HP_Contour_18 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
149 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1490330 | Import Contours | NA | 4.3.6. Temporal Enhancement Application | 1423681 | HP_Contour_19 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
150 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.3.6. Temporal Enhancement Application | 1490757 | Workflow_1 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
151 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.3.6. Temporal Enhancement Application | 1490757 | Workflow_2 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
152 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.3.6. Temporal Enhancement Application | 1490757 | Workflow_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
153 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.3.6. Temporal Enhancement Application | 1490757 | Workflow_3 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
154 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.3.6. Temporal Enhancement Application | 1490757 | Workflow_4 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
155 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.3.6. Temporal Enhancement Application | 1490757 | Workflow_5 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
156 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.3.6. Temporal Enhancement Application | 1490757 | Workflow_6 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
157 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1357906 | Workflow : Create new workflow | NA | 4.3.6. Temporal Enhancement Application | 1490757 | Workflow_7 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
158 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1234276 | Workflow: workflow selection and navigation between workflow steps | NA | 4.3.6. Temporal Enhancement Application | 1490757 | Workflow_8 | PASS | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
159 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1410580 | Perfusion fusion images | Temporal.Segmentation.Userawareness | 4.3.6. Temporal Enhancement Application | 1424524 | App_RMF_29 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None
160 | 1179490 | Suite Workflow | 1228262 | General Behavior | Section 2.2.1 and Section 2.2.2 | 1457111 | Perfusion Stress rest | Temporal.Segmentation.Userawareness | 4.3.6. Temporal Enhancement Application | 1424524 | App_RMF_29 | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820>3.3.10 | Pass | None | None

## System Keys_Traceability
SL No | URS Requirement ID | URS Feature (DHF380082,LH) | PRS Requirement ID | PRS Feature (DHF380211,Rev LJ ) | SDS ID (DHF380234 ,Rev LE) | DRS Requirement ID | DRS Feature (D000856146,Rev I) | RMM ID (D000879224 , Rev G) | DDS Id (D000993081,Rev D) | Verification Test Case ID: Ref: D001168064, Rev B | Test Case Name/Details Ref: D001168064, Rev B | Verification Test Result (Pass/Fail) [Verification_Report: D000996308, Rev E] Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | Defect ID (for Failed steps from Verification) | Validation Test ID (Validation Plan DHF380226, Rev D, Validation Report D000926820 , Rev C) | Validation Test Result (Pass/Fail) [Validation Report D000926820, Rev C] | Defect ID (for Failed steps from Validation) | Remarks
1 | 1179488 | Integration to external applications | 1228262 | General Behavior | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440072 | Different Launch Configurations | N/A | Refer Section 4.3 | 1445411 | MR Cardiac Suite-Different Launch Configurations | PASS | NA | DHF380226>12.4 | VbA | VbA | None
2 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440073 | Different ISP configurations | N/A | Refer Section 4.3 | 1445412 | MR Cardiac Suite-Master Slave Configuration | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
3 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440073 | Different ISP configurations | N/A | Refer Section 4.3 | 1445413 | MR Cardiac Suite-CDB Configuration | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
4 | 1044731 | PACS Integration | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440073 | Different ISP configurations | N/A | Refer Section 4.3 | 1445414 | MR Cardiac Suite-PACS Plugin Configuration | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | 1499816 | DHF380226 >12.4 D000926820 | Pass | None | None
5 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440073 | Different ISP configurations | N/A | Refer Section 4.3 | 1445414 | MR Cardiac Suite-PACS Plugin Configuration | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | 1499816 | DHF380226 >12.4 D000926820 | Pass | None | None
6 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440073 | Different ISP configurations | N/A | Refer Section 4.3 | 1445415 | MR Cardiac suite-Install anywhere configuration | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
7 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440073 | Different ISP configurations | N/A | Refer Section 4.3 | 1453997 | MR Cardiac Suite-Master Backup Configuration | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
8 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440076 | Logging | N/A | Refer Section 4.3 | 1445416 | MR Cardiac Suite-Logging & Reliability | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
9 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412625 | Reliability | N/A | Refer Section 4.3 | 1445416 | MR Cardiac Suite-Logging & Reliability | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
10 | 1179485 | MR Cardiac Suite, | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1445417 | MR Cardiac Suite-Monitor Support | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
11 | 691819 | Support high resolution medical grade monitors | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1445417 | MR Cardiac Suite-Monitor Support | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
12 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440086 | Hosting | N/A | Refer Section 4.3 | 1445419 | MR Cardiac Suite-Hosting | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
13 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440087 | I4 DB configurations and event handling | N/A | Refer Section 4.3 | 1445420 | MR Cardiac Suite-I4 DB configurations and event handling | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
14 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440091 | Concurrency and application weight | N/A | Refer Section 4.3 | 1445421 | MR Cardiac Suite: Concurrency and application weight | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
15 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440694 | Illumeo Configurator | N/A | Refer Section 4.3 | 1445422 | MR Cardiac Suite-Illumeo Configurator | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
16 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440089 | Upgrade Paths | N/A | Refer Section 4.3 | 1445423 | MR Cardiac Suite-Upgrade Paths | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
17 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440073 | Different ISP configurations | N/A | Refer Section 4.3 | 1445430 | MR Cardiac Suite-PIAB Configuration | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
18 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1364582 | Preprocessing Support | N/A | Refer Section 4.3 | 1445822 | MR Cardiac Suite-Pre-Processing Support | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
19 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440085 | Service Tools | N/A | Refer Section 4.3 | 1451549 | MR Cardiac Suite-Service Tools_Restartscript&PMT | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
20 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440085 | Service Tools | N/A | Refer Section 4.3 | 1451550 | MR Cardiac Suite-Service Tools_Verify E2E logs | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
21 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440085 | Service Tools | N/A | Refer Section 4.3 | 1451551 | MR Cardiac Suite-Service Tools_Logs | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
22 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440085 | Service Tools | N/A | Refer Section 4.3 | 1451552 | MR Cardiac Suite-Service Tools_RetryI4DBEvents_StopBothServices | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
23 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440085 | Service Tools | N/A | Refer Section 4.3 | 1451553 | MR Cardiac Suite-Service Tools_StudyDeletion | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
24 | 1179485 | MR Cardiac Suite | 1228262 | General Behavior | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412624 | Performance | N/A | Refer Section 4.3 | 1445425 | MR Cardiac Suite-Launch Performance | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
25 | 1179485 | MR Cardiac Suite | 1228262 | General Behavior | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412626 | Inter-operability | N/A | Refer Section 4.3 | 1445427 | MR Cardiac Suite-Inter-operability | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
26 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440090 | Securing endpoints | N/A | Refer Section 4.3 | 1445428 | MR Cardiac Suite-Security | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
27 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1412627 | Security | N/A | Refer Section 4.3 | 1445428 | MR Cardiac Suite-Security | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
28 | 1179485 | MR Cardiac Suite | 1228262 | General Behavior | Refer Section 2.2.1 Cardiac MR Suite Applications | 1438340 | Licencing | N/A | Refer Section 4.3 | 1445429 | MR Cardiac Suite-Licensing | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
29 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441362 | MR Cardiac Suite Localization-French | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
30 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441363 | MR Cardiac Suite-Localization-Danish | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
31 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441364 | MR Cardiac Suite-Localization-Germany | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
32 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441365 | MR Cardiac Suite-Localization-Italian | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
33 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441366 | MR Cardiac Suite-Localization-Sweedish | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
34 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441367 | MR Cardiac Suite-Localization-Russia | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
35 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441368 | MR Cardiac Suite-Localization-Spanish | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
36 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441369 | MR Cardiac Suite-Localization-Norwegian | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
37 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441370 | MR Cardiac Suite-Localization-Japanese | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
38 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441371 | MR Cardiac Suite-Localization-Dutch | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
39 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440074 | Localization | N/A | Refer Section 4.3 | 1441372 | MR Cardiac Suite-Localization-Chinese | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
40 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1241851 | Usage Analytics | N/A | Refer Section 4.3 | 1445416 | MR Cardiac Suite-Logging & Reliability | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
41 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237783 | Logical Series: US modality | N/A | Refer Section 4.3 | 1360522 | 1237783 - US Basic Viewing and Logical Series rules | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
42 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237784 | Logical Series: X-ray modality | N/A | Refer Section 4.3 | 1360523 | 1237784 - Non-Imaging Services - Logical Series CR&DX | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
43 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237785 | Logical Series: CT modality | N/A | Refer Section 4.3 | 1360527 | 1237785 - Non-Imaging Services - Logical Series for CT | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
44 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237786 | Logical Series: MR modality | N/A | Refer Section 4.3 | 1363426 | 1237786 - Non-Imaging Services - Logical Series for MR | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
45 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1237782 | Logical series: common behavior | N/A | Refer Section 4.3 | 1363426 | 1237786 - Non-Imaging Services - Logical Series for MR | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
46 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440092 | Copy Table to clipboard | N/A | Refer Section 4.3 | 1490328 | Copy Table to clipboard | Pass Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A | NA | DHF380226 >12.4 D000926820 | Pass | None | None
47 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1288854 | Displaying multiple applications side by side | N/A | Refer Section 4.3 | 1490493 | Multiple Applications side by side | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
48 | 1179490 | Suite Workflow | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1288854 | Displaying multiple applications side by side | N/A | Refer Section 4.3 | 1490493 | Multiple Applications side by side | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
49 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1288859 | Opening new applications | N/A | Refer Section 4.3 | 1490493 | Multiple Applications side by side | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
50 | 1179490 | Suite Workflow | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1288859 | Opening new applications | N/A | Refer Section 4.3 | 1490493 | Multiple Applications side by side | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
51 | 1179485 | MR Cardiac Suite, | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490539 | Portrait Hangings_FucntionalLVRV | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
52 | 691819 | Support high resolution medical grade monitors | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490539 | Portrait Hangings_FucntionalLVRV | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
53 | 1179485 | MR Cardiac Suite, | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490648 | Portrait Hanging-Spatial Enhancement | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
54 | 691819 | Support high resolution medical grade monitors | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490648 | Portrait Hanging-Spatial Enhancement | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
55 | 1179485 | MR Cardiac Suite, | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490649 | Portrait Hangings-QFLOW | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
56 | 691819 | Support high resolution medical grade monitors | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490649 | Portrait Hangings-QFLOW | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
57 | 1179485 | MR Cardiac Suite, | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490675 | Protrait Hangings-Functional LA | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
58 | 691819 | Support high resolution medical grade monitors | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490675 | Protrait Hangings-Functional LA | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
59 | 1179485 | MR Cardiac Suite, | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490676 | Portrait Hangings-Mapping | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
60 | 691819 | Support high resolution medical grade monitors | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490676 | Portrait Hangings-Mapping | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
61 | 1179485 | MR Cardiac Suite, | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490679 | Portrait Hangings-Temporal Enhancement | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
62 | 691819 | Support high resolution medical grade monitors | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1440080 | Monitor Support | N/A | Refer Section 4.3 | 1490679 | Portrait Hangings-Temporal Enhancement | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
63 | 1179485 | MR Cardiac Suite | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1233675 | Launching external applications | N/A | Refer Section 4.3 | 1445411 | MR Cardiac Suite-Different Launch Configurations | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
64 | 1179488 | Integration to external applications | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1233675 | Launching external applications | N/A | Refer Section 4.3 | 1445411 | MR Cardiac Suite-Different Launch Configurations | PASS | NA | DHF380226 >12.4 D000926820 | Pass | None | None
65 | 1179485 | MR Cardiac Suite, | 1182564 | Export Findings | Refer Section 2.2.1 Cardiac MR Suite Applications | 1525537 | generate RTF reporting | Viewer.DiscreteReporting | Refer Section 4.3 | 1528410 | Generate report-MRC Viewer, | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, | 1182564 | Export Findings | Refer Section 2.2.1 Cardiac MR Suite Applications | 1525537 | generate RTF reporting | FCMR.DiscreteReporting | Refer Section 4.3 | 1528413 | Generate report-Functional LVRV | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, | 1182564 | Export Findings | Refer Section 2.2.1 Cardiac MR Suite Applications | 1525537 | generate RTF reporting | FuncLA.DiscreteReporting | Refer Section 4.3 | 1528414 | Generate report-Functional LA | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, | 1182564 | Export Findings | Refer Section 2.2.1 Cardiac MR Suite Applications | 1525537 | generate RTF reporting | Mapping.DiscreteReporting | Refer Section 4.3 | 1528420 | Generate report- Mapping | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, | 1182564 | Export Findings | Refer Section 2.2.1 Cardiac MR Suite Applications | 1525537 | generate RTF reporting | QFlow.DiscreteReporting | Refer Section 4.3 | 1528429 | Generate report- qFlow Analysis | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, | 1182564 | Export Findings | Refer Section 2.2.1 Cardiac MR Suite Applications | 1525537 | generate RTF reporting | Spatial.DiscreteReporting | Refer Section 4.3 | 1528432 | Generate report- Spatail Enhancemnet | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, |  |  | Refer Section 2.2.1 Cardiac MR Suite Applications |  |  |  | Refer Section 4.3 | 1528421 | Generate report- Temporal Enhancemnet | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, | 1440095 | System Behavior for Cardiac Suite | Refer Section 2.2.1 Cardiac MR Suite Applications | 1525551 | Add to running application | NA | Refer Section 4.3 | 1527613 | Add to running | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, | 1228262 | General behavior | Refer Section 2.2.1 Cardiac MR Suite Applications | 1241456 | User preferences: common behavior | NA | Refer Section 4.3 | 1528379 | User Preferences-Automatic contours | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, | 1228262 | General behavior | Refer Section 2.2.1 Cardiac MR Suite Applications | 1318000 | Save findings | NA | Refer Section 4.3 | 1535274 | MRC Suite-FD-Save Selected Findings | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, | 1228262 | General behavior | Refer Section 2.2.1 Cardiac MR Suite Applications | 1395433 | Save and launch results | NA | Refer Section 4.3 | 1490056 | Findings: Save and Launch Results | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A
| 1179485 | MR Cardiac Suite, | 1228262 | General behavior | Refer Section 2.2.1 Cardiac MR Suite Applications | 1413756 | Send to Film | NA | Refer Section 4.3 | 1528400 | MRC Suite : Send to film | PASS Ref: D001447327_AVW_NM Suite_CT ASPECTS 15.0.5 Change Request and Level Release Plan_Rev A

## Record & Template Information
Record History
Rev | Release Date | Author | Description of Changes
A | In Windchill | Mohammed Mustaq | ISP13 Validation2: Initial version of Traceability Matrix for MR Cardiac Suite for validation2 Release.
B | In Windchill | Rajashekaraiah Lokesh | ISP13 Validation2: Updated Traceability with Revision version for all DHF ID in Header of each application tab. Updated Traceability Risk-Effectiveness tab for Column F
C | In Windchill | Rajashekaraiah Lokesh | ISP13 Validation 3: Traceability Matrix updated for Temporal and CardiacMapping applications
D | In Windchill | Rajashekaraiah Lokesh | ISP13 Validation 3: Traceability Matrix updated with Verification test Results for Val3. Removed One RMM Test with ID "5.2.3 TC_Measurements Common Hazard from TraceMatrix Spatial Enchancemen tabl" which will be executed in next release
E | In Windchill | Rajashekaraiah Lokesh | AVW15 (ISP13) RFD: Traceability Matrix updated for AVW 15 (ISP13) RFD content. 1367569 – Steps from 1 to 4 and Step 9 is removed 1366755 - Steps from 1 to 9 are removed 1366785 – Steps 1 and 2 are removed 1367400- Steps from 1 to 9 are removed 1367570 – Test case is applicable only to Viewing, hence Removing from Functional LVRV, QFlow, Spatial, Temporal, Mapping applications
F | In Windchill | Rajashekaraiah Lokesh | AVW15 (ISP13) RFD: Traceability Matrix updated with Verification test Results for RFD. Updated Viewer Traceability Matrix for RFD. In Traceability Matrix Qflow tab: Added traceability for new requirement 1412026. Added System Keys Traceability Tab For All Application except viewer added Traceability for, 1. Discrete Reporting - Req ID 1440093 For All Application except viewer and Spatial Enhancement added Traceability for 1. Copy table to Clipboard - Req ID 1440092
G | In Windchill | Rajashekaraiah Lokesh | AVW15 (ISP13) RFD: Editorial changes for RFD content.Hence previous version signature still holds good.
H | In Windchill | Rajashekaraiah Lokesh | AVW15 (ISP13) RC: Updated Revision for all DHF. Qflow: Test Case: 1367609 - Qflow_HAZ_AppSpecific_15 step is updated with to cover new requirement (1300188). Test Case:1367371 Updated with DRS key "1294456" Test Case: 1367371 - CMRQflow_Workflow3_11 step is updated with to cover new requirement (1439158). Removed Requirement "1294380" as it is not applicable for RC Functional LVRV: Test Case: 1366745 - LVRV_Haz_Results_8 step is updated with to cover new requirement (1300188). Updated Row from 198 to 225 to cover FOK site issue test case Functional LALV: Test Case: 1368292 - LALV_HAZ_8 step is updated with to cover new requirement (1300188). All Mapping RMM is updated in the Cardiac Mapping traceability sheet. Viewer: Removed the DRS keys in Viewer tab as mentioned below, 1239007 1240920 1355346 1355343 1357612 1357897
I | In Windchill | Praveenkumar S B | AVW15 (ISP13) RC: -1418189-Test set removed as requirement 1235995 is removed. -1418190-Updated testcase Basic_Viewing_1 as requirement for DXR 1240129 is removed. -1418184 Test set removed “MR Cardiac Viewer -- Exam/Series border highlight” as requirement 1240026 was removed. -Removed the test set 1360508 1360492 1360493 1360494 1360495 1360502 1360503 1360504 1360508 1360509 1360515 which are not relevant and all the requirements related to hanging protocol is covered under test sets 1360491 1360506 1360510 1363360 1364264 -Removed the test set 1358162 as the 1238804 is covered under test set. -Removed the test set 1418180 as requirement 1240013 is removed. -Updated the Test Case 1418163: Added new test steps from Cine_10 to Cine_16. -Test Case 1362945 - Steps 43 to 46 is removed -Test Case 1362946 - Steps 8 to 11 is removed -Test Case 1362947 - Steps 5 to 11 is removed -Test Case 1363466 - Step 9 is removed -Test Case 1368837 - Added new step , step 13 for missing requirement 1351127 -Removed the obsolete test steps "Viewer2_Haz_15" and "Viewer2_Haz_16" from the TC ID: 1450153 -The traceablity was updated with Pass or Fail status for all the MR Cardiac Suite applications in the respective applications tab.
J | In Windchill | Praveenkumar S B | AVW15 (ISP13) LR1 : The Traceablity matrix was updated for the LR1 impacted Test Cases, for each MRC Suite applications. The following tabs were updated. Traceablity-MR Viewer : TC ID:1367400,1450153,1450157,1450158,1418142,1418145 Traceability Matrix Func LVRV:TC ID:1366589,1266630,1366682,1366717,1366718,1367960,1366742,1366745 Traceability Matrix Func LA LV : TC ID: 1367496,1367524,1368838,1368292 Traceability Matrix QFlow : TC ID: 1367185,1367187,1367190,1367371,1367608,1456493,1367609 TraceMatrix Spatial Enhancement: TC ID: 1367012,1367035,1367074,1367078 Cardiac Mapping_Traceability: TC ID: 1420421,1420422,1420423,1420424,1420425,1420426,1474095,1490254,1532339,1535580, 1535591,1539018,1539026,1420622. TemporalEnhancement_Tracebility : TC ID: 1423681,1424155,1424524
Record Review and Approval
Name Author |  | Function Author
Praveenkumar S B |  | Verification Engineer
Name Reviewer(s) |  | Function Reviewer(s) | Review Date(s) | Signature(s)
DD-MMM-YYYY
Gaurav Balani |  | Software Architect | In Windchill | In Windchill
Aashish Kumar |  | Verification Manager | In Windchill | In Windchill
Rishov Majumder |  | Clinical Specialist | In Windchill | In Windchill
Vivek Kumar J |  | Quality Engineer | In Windchill | In Windchill
Name Approver |  | Function Approver | Date | Signature
DD-MMM-YYYY
Shilpa Halliyal |  | Project Manager | In Windchill | In Windchill
End of Document