## Slide 1
AVW15, CT ASPECTS & NM Suite RC Release –
Verification Results Review
Review Meeting Date: 26-June-2023
ISP Verification Team

## Slide 2
Review Purpose
Verification Scope
Verification Build Summary
Overall Verification Status
Overall Defects Status
Verification Defects Impact on Design Changes
Expert Test Scope & Summary
Documentation Status
Overall Verification Completion Status
Summary
Agenda

## Slide 3
Review Purpose
# | What? | Description | By Whom
1 | Verification report | Consolidated report of the tests executed during verification phase | Verification Leads/SMs
2 | Overall Verification Status | Review on verification execution and status Defects found during verification Changes in requirements from verification results, System design changes , detailed requirement changes, as per Verification design review material | Global Verification Lead
3 | Verification results impact on design | Defect fixes scoped in for verification did not had any design change & did not result in any PRS, DRS, SDS, DDS updates | Verification Leads/SMs

## Slide 4
AVW15 RC – Verification Scope
MR Cardiac Suite:
Functional Test
RMM
MR Applications:
Sanity and Regression Tests (including LOBI and NQ changes)
CT Applications:
Regression Tests
BP update regression
AVA vessels tool functional test
System Features HFA:
Labeling and Regression Tests
CT Cardiac Applications:
CAD RADS 2.0 and Findings Management
CCA RFCs
CT Spectral Applications:
RMM (impact based) and Regression
sCTV and sMGOP new requirement Functionality Test
CT Viewing:
Regression test, Sanity and Upgrades
CT Viewing - regression test
NM Applications:
NM Suite Sanity Testing
NM 3rd Party Apps Sanity & Regression(Autoquant 2017.22)
MiradaXD Integration Sanity & Regression
Knowledgescape Sanity & Regression
CT, MR, NM Legacy Applications:
MR RMM (impact based)
CT, NM(including Corridor4DM 2023)
MRC and MR Regression Tests
Installation, Upgrades & Security:
Installation, upgrades and Security Regression Tests
PIAB with both C and D drive SSD
Enable Encryption option on PMT
System Features BLR:
PMT client auto upgrade Functional tests
Regression Tests
Serviceability:
IOT RTS1.0 Functional tests
PSC9.0.1-148 Regression Tests
Axeda Split package and SPLUNK Regression Tests
Regression Tests

## Slide 5
AVW15 RC – Verification Scope
Impact of Variances (if any)
None.
None
Deployments:​
AVW Demo Laptop
AVW Server HW Configuration ​
AVW Workstation ​
AVW Enterprise
Non-Functional Tests:
Performance Tests of Applications
Reliability Test
AVW Sanity: Gen3 and A2 GPU
Variances (if any)
Compatibility Tests
# | Modality | AVW15 RC
1 | CT | CT Incisive patch (couch height fix) Incisive 1.1 HB, V4.5.1.742
2 | CT | CT Incisive for Dual Energy verification per marketing request
2 | CT | CT Thor (if the preliminary datacases are available on time)
3 | ISCV | V7.0
4 | IBE | B.19
5 | iGT (IXR) | Allura 8.x.100
6 | US | More legacy cameras if the data is available
7 | Pinnacle | IOCC verification only

## Slide 6
CT ASPECTS – Verification Scope
CT ASPECTS
  Sanity
Variances (if any)
Impact of Variances (if any)
None
Not Applicable
Deployments Verified
HX Configuration (Client Server)

## Slide 7
NM Suite – Verification Scope
NM Suite
Sanity for NM Planar Suite, NM JETPack and NM QA QC Tools
Variances (if any)
Impact of Variances (if any)
None
Not Applicable
Deployments Verified
HX Configuration (Client Server)
PIAB

## Slide 8
Client/Server and PIAB Builds:
ISP.Server_15.1.0.11708.iso Sver ISO
ISP.Server_15.0.3.11738.iso
ISP.Server_15.0.3.11754.iso
ISP.Server_15.0.3.11779.iso
ISP.Server_15.0.3.11792.iso
ISP.Server_15.0.3.11821.iso
ISP.Server_15.0.3.11834.iso
ISP.Server_15.0.3.11877.iso
ISP.Server_15.0.4.11103.iso
ISP.Server_15.0.4.11117.iso
ISP.Server_15.0.4.11137.iso
ISP.Server_15.0.4.11145.iso
ISP.Server_15.0.4.11148.iso
ISP.Server_15.0.4.11170.iso
ISP.Server_15.0.4.11176.iso
ISP.Server_15.0.4.11180.iso EV ISO
AVW15, CT ASPECTS & NM Suite – Builds Released for Verification
ISO Released for Verification:
Enterprise Builds:
ISP.Concerto_15.1.0.12708.iso
ISP.Concerto_15.0.3.12738.iso
ISP.Concerto_15.0.3.12754.iso
ISP.Concerto_15.0.3.12779.iso
ISP.Concerto_15.0.3.12792.iso
ISP.Concerto_15.0.3.12821.iso
ISP.Concerto_15.0.3.12834.iso
ISP.Concerto_15.0.3.12877.iso
ISP.Concerto_15.0.4.12103.iso
ISP.Concerto_15.0.4.12117.iso
ISP.Concerto_15.0.4.12137.iso
ISP.Concerto_15.0.4.12145.iso
ISP.Concerto_15.0.4.12148.iso
ISP.Concerto_15.0.4.12170.iso
ISP.Concerto_15.0.4.12176.iso
ISP.Concerto_15.0.4.12180.iso
Demo Laptop (HP) Builds:
IntelliSpace_Portal_Demo_Laptop_OVA_V15.0.4.11170
IntelliSpace_Portal_Demo_Laptop_OVA_V15.1.0.11703
IntelliSpace_Portal_Demo_Laptop_OVA_V15.0.4.11180
Demo Data:
AVW_Demo_Data_FSE_V15.0.0.17100.iso
AVW_Demo_Data_V15.0.4.17000.iso

## Slide 9
AVW15, CT ASPECTS & NM Suite – Builds Released for Verification
ISO Released for Verification:
AVW15 RC – Final Builds Released for Verification
Refer Section 9.5 (Final Builds (Media) Released for Verification) of Design Verification Report (D000996308_Rev F).
Media Menu Builds:
AVW15(ISP13) IntelliSpace_Portal_Media_Menu_V15.0.4.11103.iso
AVW15(ISP13) IntelliSpace_Portal_Media_Menu_V15.0.4.11176.iso
AVW15(ISP13) IntelliSpace_Portal_Media_Menu_V15.0.4.11180.iso
Unified Media:
IntelliSpace_Portal_Unified_Media_IX_V15.0.4.11180
IntelliSpace_Portal_Unified_Media_V15.0.4.11180
RSD packages:
ISP.Server_15.0.4.11180.part01.zip to ISP.Server_15.0.4.11180.part15.zip
IntelliSpace_Portal_Media_Menu_V15.0.4.11180.part01.zip to IntelliSpace_Portal_Media_Menu_V15.0.4.11180.part15.zip
KnowledgeScape ISO
KnowledgeScape_1Click_15.0.4.41882.iso

## Slide 10
10
AVW15, CT ASPECTS & NM Suite RC Release – Overall Test Execution Summary
Note:
No failures found in safety tests.
Failed tests – defects on failed tests were reviewed and either fixed and verified or postponed to a different verification phase – according to defect management procedure.
Not executed tests are considered as Not Applicable in Execution Progress graph and not considered in Overall Execution graph. (Refer verification report)

## Slide 11
# | Checkpoints | AVW15 | CT ASPECTS | NM Suite
1 | All defects found during verification for the release scope have been scored and accepted. Exception: Defects submitted during the last two weeks prior to the milestone should be set as hazard yes/no (scoring is required only for Hazard = Yes). | Yes | Yes | Yes
2 | All rank 2 defects for the release scope must be resolved and closed by end of verification. Exception: a. Rank 2 defects replicated from other projects. b. Rank 2 defects which need a fix outside the organization (example: US, MR..etc) c. Rank 2 defects logged from outside R&D. | Yes | Yes | Yes
3 | All hazard / safety test cases executed for the milestone are in passed state. | Yes | Yes | Yes
4 | All Rank 1 defects and Sev-3 are resolved and closed | Yes | Yes | Yes
5 | All the resolved defects for the release scope are closed. | Yes | Yes | Yes
6 | Sanity test is completed and passed on the Final Verification build during, end of verification. | Yes | Yes | Yes
7 | All tests including the Expert tests have been performed as per the Verification plan. | Yes | Yes | Yes
8 | All the documents needed for the end of verification (including Verification Reports, Traceability matrix, RMM Update with Test Evidence and Engineering Release Notes) are reviewed and approved. | Yes | Yes | Yes
9 | Summary | Go | Go | Go
AVW15, CT ASPECTS & NM Suite RC Release - Verification Completion Status (Applications)
Done
On-Track
Gap
At Risk

## Slide 12
Source: TFS
Enhancements
Closed / Removed
1149
Open
102
Found
1249
Product Defects
All the open defects (102) have been analyzed & moved to future milestones / releases.
Refer " D001025480_AVW15, CT ASPECTS & NM Suite RFC_Product Defect Status Report_Rev C (Section 3)" for more details.
AVW15, CT ASPECTS & NM Suite RC Release – Overall Defect Status Summary
Closed / Removed 185
Open
37
Found
223
Notes: 1. Defect dump taken from TFS on 26-June-2023 @ 10:30 AM HFA time. 2. Defect dump list the defects logged for AVW15, CT ASPECTS & NM Suite RC Release
3.Total open defects count is 102, out of which 2 defects were fixed and planned to be released in LR1.
Rank | Un Ranked | 1 | 2 | 3 | 4 | Total
Accepted | 0 | 0 | 0 | 47 | 6 | 53
Closed | 0 | 1 | 324 | 488 | 23 | 836
Investigated | 0 | 0 | 0 | 7 | 0 | 7
New | 0 | 0 | 0 | 25 | 3 | 28
Planned | 0 | 0 | 0 | 4 | 1 | 5
Removed | 64 | 0 | 37 | 163 | 49 | 313
Resolved | 0 | 0 | 0 | 1 | 1 | 2
Scored | 0 | 0 | 0 | 5 | 0 | 5
Total (All) | 64 | 1 | 361 | 740 | 83 | 1249
Total Open Defects | 0 | 0 | 0 | 91 | 11 | 102
Rank | Un Ranked | 1 | 2 | 3 | 4 | Total
Accepted | 0 | 0 | 2 | 20 | 1 | 23
Closed | 0 | 0 | 21 | 93 | 4 | 118
Investigated | 0 | 0 |  |  |  | 0
New | 0 | 0 | 1 | 5 | 2 | 8
Planned | 0 | 0 |  | 1 |  | 1
Removed | 2 | 0 | 3 | 50 | 13 | 68
Resolved | 0 | 0 |  | 3 |  | 3
Scored | 0 | 0 |  | 2 |  | 2
Total (All) | 2 | 0 | 27 | 174 | 20 | 223
Total Open Enhancement | 0 | 0 | 3 | 31 | 3 | 37

## Slide 13
Sl. No. | Application Name | Verification Lead | Did RC-. Verf defect result in design change (Y/N) | DHF # if design changes | # Open Score 1 TFS Defects | # Open Score 2 TFS Defects
AVW15 |  |  |  |  |  | 
1 | MR Cardiac Suite & MR Applications | Lokesh Rajashekaraiah | N | N/A | 0 | 0
2 | CT Applications & System Features HFA | Vladimir Barbarash | N | N/A | 0 | 0
3 | CT Cardiac Applications & CT Viewing | Chandrashekhar Venkatraman | N | N/A | 0 | 0
4 | CT Spectral Applications | Karanam, Haindavi | N | N/A | 0 | 0
5 | NM Applications | Rudravarapu Kasiviswanath | N | N/A | 0 | 0
6 | CT, MR, NM Legacy Applications | Arthi Deenadayalan | N | N/A | 0 | 0
7 | Installation, Upgrades & Security | Santhosh H | N | N/A | 0 | 0
8 | System Features BLR | Kavya Vishwakumar | N | N/A | 0 | 0
9 | Serviceability | Kavya Vishwakumar | N | N/A | 0 | 0
10 | Non-Functional Tests | Kiran B | N | N/A | 0 | 0
11 | Compatibility Tests | Vladimir Barbarash Arthi Deenadayalan Kavya Vishwakumar | N | N/A | 0 | 0
CT ASPECTS |  |  |  |  |  | 
1 | CT ASPECTS | Vladimir Barbarash | N | N/A | 0 | 0
NM Suite |  |  |  |  |  | 
1 | NM Suite | Rudravarapu Kasiviswanath | N | N/A | 0 | 0
Verification Defects (AVW15, CT ASPECTS & NM Suite) : Impact on Design Changes for RC
Notes: All open score 2 defects should have rationale to justify
‘Why is it OK to release RfVD with this defect’?

## Slide 14
AVW15 & CT ASPECTS - Expert Test Results & Status Summary
Overall status :
Application | Scope | Tested by | Status
MR Cardiac Suite | MM Viewer, Results & Workflow, Functional LA LV, Functional LV RV, Spatial Enhancement, Temporal Enhancement, T1, T2, T2* Mapping and QFlow | Ruth Katz & Rishov Majumder | Passed
CT CCA | Findings management, CAD-RADS & Multi-Batch | Shuli Ganashvili | Passed
CT Calcium Automated Analysis | Automated calcium scoring results for non-gated scans for incidental findings | Suru Narayan Srikrishna | Passed
MR LOBI – CVS (FLAIR*) | FLAIR* map to support MS cases using FLAIR and SWI | Rishov Majumder | Passed
Spectral Enhancements | PRVR & sMGoP | Katz, Ruth | Passed
Spectral AVA | Spectral AVA | Ganashvili, Shuli | Passed
Discrete Results Export | Discrete Results Export | Ganashvili Shuli and Suru Narayan Srikrishna | Passed

## Slide 15
AVW15, CT ASPECTS & NM Suite – Quality Design Control : Documentation Status
Design Control | Status | Comments
Traceability Matrix |  | All traceability documents are released.
Risk Management Matrix |  | All RMM documents Released
System Malware Scan Verification report |  | Released
Design Verification Report |  | Released
Product Defect Status report |  | Released
Design review Record - Verification |  | This Document
Overall status :
Refer section 1.3 (D001045958) of the document

## Slide 16
Exit Criteria | Current Deviation/Status | Status | Comments
All defects found during verification for the release scope have been scored and accepted. Exception: Defects submitted during the last two weeks prior to the milestone should be set as hazard yes/no (scoring is required only for Hazard = Yes). | Done |  | None
All rank 2 defects for the release scope must be resolved and closed by end of verification. Exceptions: a. Rank 2 defects replicated from other projects. b. Rank 2 defects which need a fix outside the organization (example: US, MR..etc) c. Rank 2 defects logged from outside R&D. | Done |  | None
All hazard / safety test cases executed for the milestone are in passed state. | Done |  | None
All Rank 1 defects and Sev-3 are resolved and closed | Done |  | None
All the resolved defects for the release scope are closed. | Done |  | None
Sanity test is completed and passed on the Final Verification build during, end of verification. | Done |  | None
All tests including the Expert tests have been performed as per the Verification plan. | Done |  | None
All the documents needed for the end of verification (including Verification Reports, Traceability matrix, RMM Update with Test Evidence and Engineering Release Notes) are reviewed and approved. | Done |  | None
Summary | Done |  | None
AVW15, CT ASPECTS & NM Suite – Overall Verification Completion Status

## Slide 17
All planned Tests execution has been completed.
All fixed defects in ‘RC’ were verified.
Failed tests – defects on failed tests were reviewed and either fixed and verified or postponed to a different verification phase – according to defect management procedure.
Failed tests with defects postponed to future releases are maintained in Fail status with Rationale.
Summary