| Problem Summary |  |
| --- | --- |
| PR #: TW # 4200442 - Incorrect field value is captured in Installation report. |  |
| Short Problem Statement |  |
| UDI (the GTIN part) in the installation report is different than the actual UDI as appears in the about screen for ISP / AVW product. |  |
| Detailed Problem Description |  |
| Due to a SW defect (the logic which was used to read the different fields from within the “version.js” file) that occurred in v12.1.5, the UDI number is not updated correctly in the ISP installation report. UDI (the GTIN part) in the installation report is different than the actual UDI as appears in the about screen for ISP / AVW product. All installed ISP (IntelliSpace Portal) systems in the field with v12.1.5 and onwards for example v12.1.8, 12.1.9 and AVW15 (Advanced Visualization Workspace) etc. have this issue in the installation report. Scope, Specific Versions & frequency of the problem in ISP / AVW Installation Report: Version; ISP SW 1-Click SN; UDI (the GTIN part). 12.1.4; Correct; Correct. 12.1.5; Correct; In-correct. 12.1.6; Correct; In-correct. 12.1.7; Correct; In-correct. 12.1.8; Correct; In-correct. 12.1.9; Correct; In-correct. Version; ISP SW 1-Click SN; UDI (the GTIN part). 15.0.0; Correct; In-correct. 15.0.4; correct; In-correct. 15.0.5; In-correct; In-correct. This issue was found in-house after the installation of ISP12.1.8 setup by Service team while verifying the details of Installation report. Refer defect ID -> 1543226: Incorrect UDI number in ISP installation report. Potential consequences – No potential consequences were identified to the patient/ user/ operator. - UDI is a unique number that gets generated for a release of the product, which do not have any impact on patient/ user. - The Installation report is not the primary source to check the unique identification of the product, instead About -> Help page is the primary source of information about UDI. Hence, there is no impact on operator. - There are consequences for compliance. CAPA to investigate how this issue was missed in different versions of ISP / AVW releases. Note: AVW (Advance Visualization Workspace) v15 is previously named as ISP (IntelliSpace Portal) up to v12 release. |
| Problem scope |  |
| System Element | Design Controls (QMS) |
| Horizontal analysis | This issue was first detected in-house in ISP12.1.8 installation report. Later a horizontal analysis was performed in the previously released versions of ISP12 and all the released versions of AVW 15. The results are documented in the ‘Detailed Problem Description’ section above. Checked with Radiology Informatics (RI) business unit (BU) as the same QMS is being used by this team also and found that they do have similar installation record but data field for UDI is entered manually hence, no impact of this issue on RI BU. |
| Product scope | Version; ISP SW 1-Click SN; UDI (the GTIN part). 12.1.4; Correct; Correct. 12.1.5; Correct; In-correct. 12.1.6; Correct; In-correct. 12.1.7; Correct; In-correct. 12.1.8; Correct; In-correct. 12.1.9; Correct; In-correct. Version; ISP SW 1-Click SN; UDI (the GTIN part). 15.0.0; Correct; In-correct. 15.0.4; correct; In-correct. 15.0.5; In-correct; In-correct. |
| Process scope | 1. Design & Development Procedure – 2005000905 (ICAP-P-030024) 2. Design Verification Procedure – 2005000901 (ICAP-P-030010) 3. Service Innovation Procedure – 2005000767 (ICAP-P-060002) |

Product Risk Assessment

Product safety/health risk: No risk as the issue is related to wrong UDI (the GTIN part) in installation report only. Note: UDI (the GTIN part) in the installation report is different than the actual UDI as appears in the about screen for ISP / AVW product. Product security: None

Process Risk / Other Risk Considerations –

Process Risk: No risk as the issue is related to wrong UDI (the GTIN part) information in the installation report only. This is because of the logic which was used to read the different fields from within the “version.js” file. No change in the process documents required. Compliance/Regulatory Risk: Yes - Severity level of impact on the Quality System (High/Medium/Low/None): Low – no impact on the product. Regression issue, missed to test the scenario in Verification during Product Development. - Probability (issue will result in detriment to the Quality System) level (High/Medium/Low/None): Medium, as mentioned in the product scope section above. - Description and justification of Severity and Probability levels: According to the review analysis done with the results as mentioned above in product scope section; the severity is low and probability of occurrence is medium. - [Risk Level (Critical/Major/Minor/None)]: Minor as there is no impact to the product or customer.

| Containment - Product |  | Action Owner | Date in Place |
| --- | --- | --- | --- |
| CT# | We identified no containment action is needed as there is no impact on the product/customer. Note: ISP SW 1-Click SN available in the installation report helps to uniquely identify the installed product. | Not Applicable | Not Applicable |

| Containment – Non-Product |  | Action Owner | Date in Place |
| --- | --- | --- | --- |
| CT# | None | Not Applicable | Not Applicable |

| Correction(s): |  | Action Owner | Date Initiated / Completed |
| --- | --- | --- | --- |
| CR# | Scenarios with above issues identified during the investigation shall be fixed and verified in ISP12.1.10 & AVW15 LR2 and documented as part of ISP12.1.10 & AVW15 LR2 report respectively. ISP12.1.10: To Be Updated AVW15 LR2: To Be Updated | Arthi Deenadayalan – Verification Lead | 16 NOV 2023 / 30 MAY 2024 |

| Communication to (potentially) impacted stakeholders: |  | Action Owner | Date Complete |
| --- | --- | --- | --- |
| CM# | Refer the attached email. | Aashish Kumar – Verification Manager | 20 FEB 2024 |
| Communication to other Comp. Domains |  | Action Owner | Date Complete |
|  | Communication to other compliance domains (like finance, privacy and legal) is not applicable as the issue is limited to the scope of ICAP internal project management and design and development domains. | Not Applicable | Not Applicable |

Root Cause – Proposed Action Plan: look for root causes for every Failure mode (if not a factor list justification)

Key Comments about Root Cause Investigation:

Problem Statement: UDI (the GTIN part) in the installation report is different than the actual UDI as appears in the about screen for ISP / AVW product. RCA Method – 5 WHY Why did this issue occur? - 1st Why – In the installation report the GTIN part of UDI is reading the older version value. - 2nd Why – Logic & sequence which are used to read different fields from within the “version.js” file is faulty. - 3rd Why – Logic & Sequence is defined to read a specific format & length of UDI number. - 4th Why – Detailed requirement is missing. [RCA1] Note: Installation Report Template contains detailed information which is not part of the requirement. - 5th Why – No reference in the requirement to use the details from the Installation Report Template (ID: 2005000819 / ICAP-T-060002.06_Installation_Report_Template). [RCA2] Why not found in Product Development of multiple releases? - 1st Why – In the installation report the GTIN part of UDI is reading the older version value. - 2nd Why – Regression issue as it was working fine till ISP12.1.4. - 3rd Why – A change was introduced in 12.4 CT ASPECTS release which causes this regression issue. - 4th Why – Detailed requirement is missing. [RCA1]

| What Root Cause Investigation Tools did you use? (Ensure copies / photos of the work are kept and attached to the Trackwise record) |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Fish Bone | PRIDE/ Pareto | Design of Experiments | Fault Tree Analysis |  | Is / Is Not | Event Sequencing |  | Cause / Effect | Contradiction Matrix | 5 WHY |  | Other: ______________ |  |
| Conclusions of Root Cause Investigation and Action Plan |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Root Cause Conclusion(s) |  | Root Cause Verification / Evidence (How has it been confirmed?) |  | Corrective / Preventive Action to Mitigate Root Cause |  |  | Action Verification and/or Validation Plan (Prove does no harm before implemented) |  |  |  | Date Action to be Implemented |  | Action Owner |
| [RCA1] – Detailed requirement is missing. [RCA2] – No reference in the requirement to use the details from the Installation Report Template (ID: 2005000819/ICAP-T-060002.06_Installation_Report_Template). |  | Refer above section “Key Comments about Root Cause Investigation: “. |  | Correction: Issue to be fixed in the latest ISP12 & AVW15 releases. |  |  | - Verification Evidence (ISP12.1.10): D001496622 – To Be Updated. - Verification Evidence (AVW15 LR2): D001573771 – To Be Updated. |  |  |  | - ISP12.1.10 – 30 MAY 2024 - AVW15 LR2 – 30 MAY 2024 |  | Arthi Deenadayalan – Verification Lead |
|  |  |  |  | Corrective Action: - Update AVW15 PRS document with the installation report to be as mentioned in the Installation Report template. - A new test case verifying the Installation Report details to be designed and executed. |  |  | - PRS Document Update: DHF380211 – To Be Updated. - Test Case Design: D001507600 ISP12.1.10_Test Specification -> Section 5 attachment (Test case id: 1546829). - Test Case Execution (ISP12.1.10): D001496622 – To Be Updated. - Test Case Execution (AVW15 LR2): D001573771 – To Be Updated. |  |  |  | - PRS Update – 30 MAY 2024 - ISP12.1.10 – 30 MAY 2024 - AVW15 LR2 – 30 MAY 2024 |  | PRS Update: Nisha Kamaluddin – Project Manager Test Case Execution: Arthi Deenadayalan – Verification Lead |
|  |  |  |  | Preventive Action: Update Service Innovation Procedure (ID: 2005000767) to include the information available in Installation Report Template in the PRS document. |  |  | Release of updated Service Innovation Procedure (ID: 2005000767) |  |  |  | 30 MAY 2024 |  | Chitrai Senthamil Selvam – Service Innovation Manager |

| Effectiveness Check Criteria – How will you verify that the Corrective/Preventive Actions are effective at preventing the defect from re-occurring? |  |
| --- | --- |
| What will be reviewed? | ISP12.1.10 & AVW15 LR2 installation Reports will be reviewed from Customer site. |
| What is the sampling plan including the rationale for the sampling decision | The first 10 installation reports of ISP12.1.10 and AVW15 LR2 product on the customer site. UDI is the unique number that gets generated from the software for a release so there is no chance of getting different UDI number for the same release. Hence, if wrong UDI is populated in the installation report, it will be found in the first sample itself. Still, to avoid any gaps a total of 20 installations report will be reviewed. |
| What is the acceptance criteria? | The acceptance will be of 100% i.e, all 10 installations of ISP12.1.10 and all 10 installations of AVW15 LR2 – the UDI in the installation report is same as in the about screen for ISP / AVW product. |
| Who will be conducting this review & documenting the results | Faina Abramovich, Chitrai Senthamil Selvam & Arthi Deenadayalan. |

| Expected Timeline: | Implementation Complete: (ambition: 90 days from Request Start) <01 SEP 2024> | CAPA Closure Complete: (ambition: 180 days from Request Start) <15 OCT 2024> |
| --- | --- | --- |