- Dicom uses TCP connection (transport connection). Well known port is 104.

The DICOM Network Protocol architecture looks something like this: Network ⇒ TCP/IP ⇒ DICOM ULP for TCP/IP ⇒ UL    Service boundary ⇒ DICOM Message Exchange ⇒ Medical Imaging Application

- Dicom is a protocol dealing with medical images.

- DICOM files typically have .dcm extension and data contains both patient data and the image/pixel data. The patient data comes from the EMR/EHR/HIS systems as HL7 data which gets tightly coupled with the equipment, procedures and image/pixel data created by the radiology medical imaging devices as DICOM data.

- The DICOM protocol is a binary Upper Level Protocol (ULP) over TCP/IP. Well known ports used by DICOM are 104, 2761, 2762 and 11112

Port 104 – PORT 104 makes possible the transmission of a datagram message from one computer to an application running in another computer.

Port 2761 - TCP port 2761 uses the Transmission Control Protocol. TCP is one of the main protocols in TCP/IP networks

Port 2762 – registered for DICOM using Transport Layer Security (TLS) over TCP or UDP.

Port 11112 - registered for DICOM using standard, open communication ...

It is used to process DICOM data, transmit, search/query, integrate, distribute, print, share, store, display medical images and patient data from the radiology archival/storage systems (PACS, RIS) to the workstation for the Radiologist to write reports.

5.   DICOM, HL7 data and network are crucial to the healthcare industry

6.   The basic DICOM communication protocols for query/retrieval, storage and print classes.

7.  The accession number is usually assigned by the HIS/RIS system and can be repeating or unique depending on the system.

8.  DICOM Objects are known as the Information Object Definitions (IOD). All real world data like patients, studies, medical devices, images, patient schedule list, a queue to be sent to a printer are objects with defined templates. These are definitions of the information to be exchanged between a Service Class User (SCU) - Client and Service Class Provider (SCP) - Server.

____________________________________________________________________________________________________________

- A DICOM modality is a property/attribute of the DICOM data object, e.g. CT, MRI, X-rays etc. are the modalities

- A DICOM Message is composed of a Command Set followed by a conditional Data Set.

- Command and Data sets are made of Elements.

- Elements have a Tag, length and value.

- Tags have group tag and element tag.

___________________________________________________________________________________________________________

- DICOM applications provide the services required for the data exchange.

DICOM main services are:

- Query/retrieve

- Print

- Store

- Storage commitment

- Worklist

DICOM Service Element (DIMSE) are used by Application Entities (AE). DIMSE Stands for DICOM Message Service Element and is the standard method through which DICOM archives communicate. We support this messaging standard for the retrieval of study, series, and instance metadata because it is widely support. For certain PACS systems, it also (currently) provides faster query results than DICOMWeb. There are two types of DIMSE Services, Composite DIMSE-C and Normalized DIMSE-N.  Five DICOM operations that fall under the DIMSE-C service:, find, move, store, get and echo.

(DICOM files contain metadata that provide information about the image data, such as the size, dimensions, bit depth, modality used to create the data, and equipment settings used to capture the image.)

These support operations and notifications like storage, retrieval, printing etc. on the SOP instances.

- AEs have Titles (AET). AE has an IP address assigned. AET are case sensitive and unique.

- Service Object Pairs (SOPs) have an unique ID (UID). SOP Classes are the fundamental unit of DICOM interoperability.

- 4 stages of DICOM – Patient, study, series and Instance.

- The DICOM Message Service Element (DIMSE) provides the mechanism for message exchange between peer Application Entities once an Association between them has been established.

- DIMSE-N provides a single Notification Service, the N-EVENT-REPORT. The N-EVENT-REPORT service is invoked by a DIMSE-service-user to report an event about a SOP Instance to a peer DIMSE-service-user. This service is a confirmed service and a response is expected.

_________________________________________________________________________________________________________

How does the DICOM transaction happen?

- AE of SCU - Client uses DIMSE to negotiate a SOP Class with the AE of SCP - Server using it's DIMSE. There are DIMSE protocols defining the procedures and encoding rules to construct messages.

- Message transactions between the two devices using DICOM begin with an Association establishment. Both devices negotiate the information structures that will be exchanged, the services that will be invoked, byte order and data compression method.

- DICOM ULP consists of seven Protocol Data Units (PDUs). Each PDU has a maximum length. PDUs are the message formats exchanged between peer entities within a layer, based on the request and the response DICOM messages.

_________________________________________________________________________________________________________