# HL7 Integration

## Patient Information Reconciliation

RIS

HIS

HL7 Component

ADT feeds ADT^A04, A08 etc

Trigger updates

Reconcile Study Info

DB

ISP

- It shall be possible for Portal to listen to the Patient Demographic changes/updates from incoming HL7 ADT messages and make the respective updates/corrections in the Patient Database. The updates shall make sure that the old demographics are preserved. Need to define the list of old attributes to be retained for comparison.

- This is achieved via an HL7 Component (IBE or something similar), which shall listen to incoming ADT Messages and trigger Portal, whenever there are relevant updates for Patients existing in Portal.

- It shall be possible to configure the list of ADT message types which shall trigger reconciliation. By default all ADT messages shall trigger reconciliation

- The update shall happen for the patient records existing in any of the local folders associated with the Portal. The list of folders which shall enable Reconciliation can be configured.

- When a patient record has been Reconciled, any of the access to this patient information, shall reflect the latest updated information. Following are some examples:

  - Display of Patient Demographics in PD

  - Viewing in Applications and the Annotations on the viewers (Image Titles)

  - Export of study to another DICOM node

  - Sending study to report. Existing draft reports should also reflect the updated information.

  - For Approved PDF reports and SC Images, the demographic updates shall be done in the DICOM header information.

- It should be possible to Enable/Disable Reconciliation via the Preference. This also include finer aspects like

  - Whether to trigger reconciliation for Open studies, and ask the user to reload demographics

  - Fetch priors for Reconciled data

  - List of tags to reconcile.

  - Separate settings needed for PACS & RIS re-conciliation?

- Enhance Prefetch to send separate queries for old and new Patient Info (if there are changes to ID & name)

### Use cases to be considered for PIR

Case 1: Patient Update is sent after the portal receives the study from Scanner (before the study is read).

ISP Action:

- ISP database gets updated to reflect the new demographics

- If Prefetch is already triggered yielding nothing, then a new Prefetch should be triggered for reconciled information.  If the Prefetch has yielded some results (i.e. Priors exist for this study, then the study should be flagged, without updating the reconciled information).

- If the data is already synch’ed to destination Portal server, should this update be communicated?  If the reverse synch of the results reconcile with the updated patient on the source, then reconciliation is not required to be propagated to the Sync Node.

Case 2: Patient update is sent after acquiring MWL from scanner. Portal receives images from scanner containing the old name.

ISP Action: Similar to Case 1.

Case 3: Patient Update is sent after the portal receives the study from Scanner and the case is already open in the viewer.

ISP Action:

- ISP database gets updated to reflect the new demographics.

- Should the user be notified and asked to reload the demographics (Preferences)

- If Prefetch is already triggered yielding nothing, then a new Prefetch should be triggered for reconciled information.  If the Prefetch has yielded some results (i.e. Priors exist for this study, then the study should be flagged, without updating the reconciled information).

- If the data is already synch’ed to destination Portal server, should this update be communicated?  If the reverse synch of the results reconcile with the updated patient on the source, then reconciliation is not required to be propagated to the Sync Node.

Case 4: Patient update is sent after the portal receives the study and the study is already Read by the Physician and the results are not yet archived.

ISP Action:

- ISP database gets updated to reflect the new demographics.

- While exporting to the PACS newly updated information is sent.

- If a draft report has been created, then the demographics get updated to the new information.

- For approved PDF and for approved SC Image reports, the DICOM header is modified.

Case 5: Patient update is sent after the portal receives the study and the study is already Read by the Physician and the results are archived.

ISP Action: Same as Case 5. We rely on the PACS to update the results as well.

Case 6: Patient update is sent after the portal receives the study and the study is already Read by the Intern and the results are generate, which need to be Reviewed & Approved.

ISP Action: Same as Case 5.

### PIR vs. Current ISP Use cases

| Use case | Impact |
| --- | --- |
| Import Study Network/Media | 1. Reconcile. Wait for Study Completion New data coming with old demographics – Reconcile. |
| Display in PD | Reconcile data and refresh PD atomically |
| Loading in Viewer/App | Notify the user that there is a Patient update. Please reload the image. Refresh the screen with the new annotations |
| Send to Film | Send display/image to film as is |
| Send to Report | Send to report as seen in viewer. If it is updated then updated demographics are sent, if not then they stay stale. |
| Print | If the data is already sent to film (as Image or display) and the image is reconciled, 1. Then a warning is shown to the user to refresh the view in case of send Image to film 2. Warning to recreate the film, in-case of send display. |
| Approve Report | If the data is already sent to report as Image and the image is reconciled, 1. Then a warning is shown to the user to refresh the demographics in-case no image is sent to report. 2. Warning to recreate the images in the report, in-case report contains images. |
| Save Result | Reconcile if the parent is reconciled Choosing PACS/Remote device on save, should honour the preferences. |
| Export Studies/Media | Always reconcile based on Preferences. If the export is in progress then don’t apply reconciliation for the data being exported. |
| Export Report | Always reconcile the patient info |
| Export Findings | Always reconcile the patient info |
| Prefetch | Prefetch using old and new demographics |
| Data Synch | Do not update synch data. If it’s already done or in progress for a study |
| Pre-processing | Pre-processed results should be reconciled |
| Reverse Synch | The reverse synch data upon receival should be reconciled. |
| CPD | Reconciliation should not impact the process |
| Anonymize | Reconciliation should not impact the process |