Application Requirements Specification

For

System Components in Philips IntelliSpace Portal 5.0

| Document Identification | TBD |  |  | PRP Reference No: |  | XCTW-0301001 Para. 8.3.2.13 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Document Version | TBD |  |  | Artemis Code |  |  | 3E.18 |  |
| Date: 04 August 2011 |  |  |  |  |  |  |  |  |  |

|  | Name | Signature | Date |
| --- | --- | --- | --- |
| Prepared by: | Dr. Kumara Sanjaya |  |  |
| Reviewed by: | Elran Gil |  |  |
| Approved by: | Noa Zoller |  |  |
| Approved by: | Nir Berenzweig |  |  |

Table of Contents

1	Introduction	6

1.1	Application Overview	6

1.2	Clinical Background and Purpose	6

1.3	Product Scope	6

1.4	Application Main Functions (tracked to PRS IDs)	8

1.5	User Characteristics	10

1.6	Assumptions and Dependencies	10

1.7	References	10

1.7.1	Internal Process Documents	10

1.7.2	Standards Documents	10

2	Requirements	11

2.1	Dual Monitor Support and Span Mode (ISPV5_PRS_ID_1_7)	11

This is not a new requirement for ISP 5.0. ISP 4.0 also supported dual monitor mode and SPREAD MODE. The following requirements are written to clarify certain additional aspects over the ISP 4.0 requirements. The ISP 4.0 System ARS should be used as a baseline for the fundamental requirements in this regard.	11

2.1.1	System will support spanning over 2 screens. This option will be available whenever 2 or more screens are connected to the client (not be protected by license key). The Span mode is also known as SPREAD MODE and has been supported in ISP 4.0 also. The only addition to this requirement in ISP 5.0 is that MMV will support SPREAD MODE and NM Comprehensive Review application will support SPREAD MODE. No other MR Applications or NM applications including 3rd party NM Applications will support SPREAD MODE. All CT applications including review applications will continue to support SPREAD MODE as they did in ISP 4.0.	11

2.1.2	The dual monitor support mode is not a new requirement and is the same as was defined for ISP 4.0. The System ARS for ISP 4.0 describes in detail about the dual monitor support requirements. All applications including CT, MR and Philips NM applications will support dual monitor configurations. However, NM 3rd party applications do not support dual monitor.	11

2.1.3	For PIB (Portal in a Box), both monitors in dual monitor mode will have the same size, aspect ratio and resolution (including DPI).	11

2.1.4	For Client-Server portal configuration, the monitors may be of different resolutions and the requirement is as detailed in the ISP 4.0 System ARS.	11

2.1.5	The process to set up dual monitors on the system will remain the same as per the implementation in ISP 4.0	11

2.2	Resizable Windows (ISPV5_PRS_ID_1_8)	11

2.2.1	In ISP 5.0 client-server configuration, all viewer and application (including CT, MR and NM except for 3rd party NM applications) windows should support the ability to be minimized, maximized and also be able to be switched across screens.	11

2.3	Multi-Resolution Support (ISPV5_PRS_ID_1_9)	11

Multiple resolution support is not a new requirement for ISP 5.0. The ISP 4.0 System ARS defines the requirements as they were implemented in ISP 4.0 Additional relevant points are provided here.	11

2.3.1	Portal Client shall support 2 resolutions (two ends of the spectrum):	11

2.3.2	The following resolutions should work without hitches:	11

2.3.3	Resolution should be determined when the application is launched. Run time adjustment to changes in resolution is not mandatory.	11

2.3.4	The System UI will adjust itself to the screen resolution at the Client Side. The adjustments will be along the following lines:	12

2.4	Support for lossy compression (ISPV5_PRS_ID_2_2)	12

2.4.1	This is not a new requirement – it is the same requirement as implemented in ISP 4.0. However the scope of support is defined here.	12

2.4.2	The system shall allow and support the following compression levels:	12

2.4.3	Compression will only affect image display.	12

2.4.4	Saved images will have no compression and therefore ensure full quality.	13

2.4.5	The system will clearly display to the user in terms of a message (status bar, header etc.) that the images being viewed are compressed images if any forms of compression is applied.	13

2.4.6	Measurements on images will always be done using the lossless data	13

2.4.7	Support for compression will be across all CT, MR and NM applications.	13

2.5	Copy Images to Client’s Clipboard (ISPV5_PRS_ID_2_6)	13

2.5.1	This is a requirement which was implemented in ISP 4.0 – but only for CT applications. The scope of this requirement is now expanded into MR Apps. (NOTE: NM Apps Scope is to be confirmed - TBD)	13

2.5.2	All portal clients will allow the user to copy images to the system clipboard.	13

2.5.3	The keyboard shortcut of [Ctrl+C] will be assigned to copy the images (selected images) to the windows clipboard.	13

2.5.4	The image will be copied using a lossless rendering into the clipboard using a common format (bmp/png etc.) and user will be able to paste it into software that supports it. In ISP5.0, only CT and MR applications will have this feature. NM Support is to be confirmed.	13

2.6	Presentation State Support (ISPV5_PRS_ID_7_4)	13

2.6.1	Presentation states shall be supported in all viewing and clinical applications – CT and MR Applications. NM Support is yet to be confirmed.	13

2.6.2	It should be possible to save presentation states.	13

2.6.3	The system should be able to load saved presentation states.	13

2.6.4	When images are sent to filming application, it should be possible to apply presentation states and send the images to filming; filming should support filming images that contain presentation states (applied presentation states).	13

2.6.5	When images are sent to reporting application, it should be possible to apply presentation states and send the images to the reporting application; reporting should support adding images containing applied presentation states into summary images and then onto the report.	13

2.7	Pre-processing and Background Processing (ISPV5_PRS_ID_7_5), (ISPV5_PRS_ID_7_18)	13

2.7.1	Pre-Processing mechanism should be supported to enable automatic data processing based on configurable fields – such as Protocol Name, Body Part, Series Description and Procedure Description.	13

2.7.2	In EX configuration, Pre-processing should be able to run on both the master and slave servers.	13

2.7.3	Manual Configuration of Pre-Processing: This should be possible within preferences and the following parameters are supported:	13

2.7.4	When a field is empty, the system will not use the field in order to choose proper series	13

2.7.5	It shall not be allowed to have all fields empty	14

2.7.6	For all test fields the system will test whether the defined value is contained in the DICOM field.	14

2.7.7	System will not look for exact match and will disregard upper/lower case	14

2.7.8	Algorithm name shall display a pull-down of all available preprocessing algorithms	14

2.7.9	Pre-processing will work if the “Pre-processing License” is valid.	14

2.7.10	UI mockup of table	14

2.7.11	Automatic Configuration of Pre-Processing: When user that has permission to change system wide settings runs a manual preprocessing, a dialog will open in PD asking whether to add the series to the automatic preprocessing list. If the user approves, then the algorithm will be added with all the DICOM fields filled according to the series. {The Scope of this requirement is yet to be confirmed}	14

1.	Support pre-processing algorithms that receive multiple series as an input either from the same study or from multiple studies (belonging to the same patient) and create one or more result series as the output (Liver, Tumor Tracking). Currently these applications do not support multiple series inputs. Hence this requirement should be considered for future versions of IntelliSpace Portal.	14

2.8	Launch Applications from another application (ISPV5_PRS_ID_7_12)	14

2.8.1	This is not an entirely new requirement. A part of this requirement with reference to CT Applications has already been implemented in ISP 4.0. The following points detail the requirements from NM perspective, which is newly added in ISP 5.0. The requirements are based on similar requirements detailed for EBW NM 2.0.	14

2.8.2	It shall be possible to launch an NM application from a running NM application. When launching, all loaded data plus all data created in the running application shall be loaded into the new application. Launching of the new application shall be done from a button in the NM applications.	14

2.8.3	The above requirement is for launching – Review Application from Analysis Application, Review Application from Review Application, Analysis Application from Analysis Application and Analysis Application from Review Application.	15

2.8.4	Launching Application from Application (LAFA) will work for CT Applications in the same way as it is implemented in ISP 4.0; click on Review  opens CTV, or switching to currently running review application; click on Analysis  open application list (including NM applications), or switching to currently running analysis application.	15

2.8.5	If launching application from application requires saving data to the disk, this data shall be saved into a temporary folder and not to a folder visible to the user. The data in the temporary folder shall be periodically deleted in order to free disk space.	15

2.9	Web based training (On Demand) (ISPV5_PRS_ID_7_25) (TBD) – will be done shortly	15

2.10	Loading multiple studies into an application (ISPV5_PRS_ID_7_27)	15

2.10.1	The system shall allow multiple datasets (studies) to be loaded into an application.	15

2.10.2	The applications will make use of this feature to optimize workflow within the application.	15

2.10.3	In NM Applications, the following behavior is expected: The selections of data to be launched into the application supported are:	15

2.11	Languages and Localization (ISPV5_PRS_ID_9_1 to ISPV5_PRS_ID_9_3)	15

2.11.1	All clinical applications and viewers shall be localized (GUI) to the following language: (ISPV5_PRS_ID_9_1)	15

2.11.2	The exceptions to the localization are:	15

2.11.3	DICOM localization will be covered in detail in the Connectivity ARS. (ISPV5_PRS_ID_9_2)	15

2.11.4	Localization of Report Templates ((ISPV5_PRS_ID_9_3)	15

2.12	Save/Load Bookmarks support for clinical and viewing applications (ISPV5_PRS_ID_2_1, ISPV5_PRS_ID_7_10)	16

2.12.1	All CT applications and viewers in IntelliSpace Portal 5.0 will support saving/loading of application state using the concept of bookmarks.	16

2.12.2	Bookmarks will not be supported by MR Applications (Including MMV) and NM Applications.	16

2.12.3	Bookmarks are “pointers” to the state of an application. In any application or viewer (except the ones mentioned earlier), it should be possible to save the state of an application at any given point in time as a “bookmark”.	16

2.12.4	The application should save the entire state in which the images/studies are at that point in time including any results created therein.	16

2.12.5	It should be possible to load the application and restore it to the state it was when the bookmark was created by selecting the saved bookmark at any point of time later.	16

2.12.6	Bookmarks should be accessible from the Patient Directory.	16

2.12.7	A single bookmark can comprise of a number of “states” within it. When the bookmark is viewed in the patient directory window, it should be possible to expand the bookmark to see the states (individual states) contained within the bookmark. When the user selects a state, the user should be able to load that state. The application which was used to create the bookmark should be launched and the state opened.	16

2.12.8	Bookmarks should have generic information which can be shared across applications (what exactly constitutes this information needs to be discussed and detailed).	16

2.12.9	The user should be able to delete bookmarks either collectively as all states within a given bookmark or individual states within a given bookmark. The system will however, allow the user to delete one bookmark at a time.	16

2.13	Data Prefetch (ISPV5_PRS_ID_7_28)	16

2.14	PACS Integration (ISPV5_PRS_ID_7_21), Desktop integration with iSITE PACS (ISPV5_PRS_ID_7_22) and Desktop Integration with 3rd Party PACS Vendors (ISPV5_PRS_ID_7_23) [author – Kishan Harwalkar]	18

2.14.1	Portal shall have API based integration with iSite PACS	18

2.14.2	It shall be possible to launch portal applications on iSite by selecting one or more series/study.	18

2.14.3	User shall be able to select the series to be launched via the series selector when not all series are available on iSite.	19

2.14.4	It shall be possible to launch the bookmarks (created by Portal) stored on iSite server. The iSite Client shall directly launch the Portal application which created the bookmark.	19

2.14.5	The bookmark shall be launched with the right portal application irrespective of whether the referenced series are stored in iSite or Portal	19

2.14.6	Fast data storage from iSite to Portal.	19

2.14.7	It shall be possible for the user to install Portal Client during the iSite Client installation.	19

2.14.8	iSite Authentication	19

2.14.9	CDP (TBD)	19

2.14.9.1	Portal shall have API based integration with CDP PACS.	19

2.14.9.2	User shall be able to select the series to be launched via the series selector when not all series are available on CDP.	19

2.14.10	3rd Party PACS Integration:	19

2.14.10.1	API based tight integration shall be provided for the following 3rd party PACS (TBD)	19

2.14.11	Utilize 3 monitors: 1 – PD & report ; 2 – Review ; 3 – Analysis ; Film can be displayed in 2 or 3 (user settings).	19

2.14.12	Automatic storage of results to PACS.	19

2.14.12.1	When portal applications are launched by selecting Study/Series from PACS Worklist (via the PACS Plug-in), the results generated shall be automatically stored on the PACS.	19

2.14.12.2	The results shall be stored on Local folder of Portal as well, if the selected data existed only in the local folder of portal.	19

2.14.12.3	It shall be possible to allow filming and printing using Portal filming application with the PACS Integration. Printing shall be done on to the printer configured in the Portal and not on the PACS.	19

2.14.12.4	Filming shall be disabled in the iSite Integration	19

2.15	Portal in a Box (Equivalent to Single Box Workstation)	19

2.15.1	Concept of Portal in a Box (PIB)	19

2.15.2	Intended clinical use of PIB:	20

2.15.3	Requirements where PIB is similar to client-server mode of ISP 5.0:	20

2.15.4	It will not be possible to Minimize and Maximize the Windows running in PIB:	20

2.15.5	Automatic logout from Windows Operating System session on exiting PIB workflow (for non-admin clinical users):	20

2.15.6	Login into PIB:	20

2.15.7	License Check on PIB Launch:	21

2.15.8	Data Download Scenarios in PIB:	21

2.15.9	Help Window in PIB:	21

2.15.10	Display of Logged in User information on the Main Screen of PIB:	21

2.15.11	Pre-configured users in PIB:	22

2.15.12	Changes in Preferences setting – available options for Admin and Non-Admin users in PIB:	22

2.15.13	User Management in PIB:	22

3	Other Nonfunctional Requirements	22

3.1	Users Rules	22

3.2	Safety Requirements	22

3.3	Database	22

3.4	Regulatory	22

3.5	Service s and Manufacturing	23

3.6	Site Adaptations	23

3.7	Event Logging	23

4	PRS/PMR Requirements   to   Use Cases/Requirements   Tractability	23

4.1.1	23

5	Revision History	23

Appendix A: Glossary	23

Appendix B: To Be Determined List	23

# Introduction

## Application Overview

## Clinical Background and Purpose

IntelliSpace Portal is a client server solution that provides advanced visualization tools for reading of CT, MR and NM images. It serves within Philips as a replacement to individual workstations that were developed in CT, MR and NM business units.

System component ARS summarizes the improvements and changes planned for the system aspects of Portal 5.0. Its purpose is to allow effort estimation for changes in the system level.

Requirements which will be addressed in the next version of ISP (6.0) will be included in an appendix at the end of the document.

## Product Scope

Philips Workspace Portal is a thin-client applications-server that provides distributed access to workstation functionality virtually anywhere the user wants. This novel solution has been hugely successful in certain areas, even with limited 3D functionality and no dedicated clinical applications to date. As we advance through the first decade of the 21st century, it is essential that analysis work spots with clinical applications are made available anywhere the clinician desires if workflow and productivity are to be optimized.

Advances in imaging technologies for CT, MR and NM are making advanced visualization a necessity, not merely an option. Technology is rapidly evolving on both the acquisition and post-processing side of the clinical equation and as clinical applications become an everyday routine instead of an exception, these applications become a critical component for the customer to differentiate the purchase decision and prove comparative-effectiveness. Beyond the specific clinical possibilities brought by specific applications, workflow considerations are crucial in the workstation market, specifically the ease-of-use and ‘access anywhere’, including the use of a centralized database, aspects of Medical Imaging Workstations.

Thin-client server technology provides widespread access to a large number of work spots, which themselves have limited individual processing power and software capabilities, is now a common method to deploy 3-D and clinical applications solutions. Therefore the thin-client model enables enterprise-wide deployment of clinical applications away from the traditional high-power radiology work spots. Specialty companies such as Tera Recon, Vital Images and Visage push the innovation threshold and lead the market toward a “simple” solution that follows the user (enterprise thin-client).

IntelliSpace Workspace Portal is a smart solution for medical professionals that bring rich CT, MR, and NM applications with actionable information and foster clinical collaboration to improve outcomes. The Portal is based on Brilliance Workspace; the world’s largest installed base of award winning is a thin-client applications-server and provides distributed access to their workstation functionality virtually anywhere the user wants.

IntelliSpace Portal 5.0 release intended to replace v4.0 in forward production systems and be available as an upgrade to the installed base of IntelliSpace Portal systems.  Based on client-server architecture, IntelliSpace Portal, Philips' multi-modality enterprise solution for CT, MR, and NM scanners, goes beyond the limits of current traditional workstation s by providing enterprise access to advanced clinical application. Using a graphical user interface (GUI) , IntelliSpace Portal's usage is efficient, intuitive, and minimizes training for customers.  IntelliSpace Portal provides advanced clinical applications anywhere within the enterprise the clinician requires. With a simple installation of software on a remote client PC, a user can access studies nearly anywhere in seconds for image interpretation or consultation.

IntelliSpace Portal 5.0 is designed to replace the current extended Brilliance Workspace (EBW) and EBW-NM workstations as the stand-alone workstation configuration for Philips CT/NM customers as well as replace the current enterprise Brilliance Portal 2.6 and IntelliSpace Portal v4.0.

IntelliSpace Portal v5.0 will be the first portal release to offer the complete NM workspace functionality.  It will bring the NM functionality realized in EBW NM 2.0 to IntelliSpace portal and make further enhancements to NM review and workflow capabilities.

IntelliSpace Portal v5.0 will enable NM workspaces to be based upon on a common platform offered either as enterprise solutions or as a “portal in a box” workstation configuration. Thus IntelliSpace Portal v5.0 will move NM workspaces to a common platform for workstation and as well enterprise solutions

The new product version includes important functionality to address competitive gaps in the current product offering:

- A rich set of CT/NM/MR clinical applications over client server

- Support a standalone workstation deployment scenario as replacement for current EBW v4.5 & EBW-NM v2.0 workstations

- Adding Multi-Vendor support capabilities to several MR Applications and Multi-Modality Viewer

## Application Main Functions (tracked to PRS IDs)

The following table lists the various PRS Key IDs which describe System Component Features. For each key/set of keys, the topics are given and the ARS documents in which the requirements will be fully detailed are given. Also the owners of the requirements are detailed for traceability.

| ARS Name | ARS Owner | PRS Key/s | Topic/s | Feature Owner | Feature Status |
| --- | --- | --- | --- | --- | --- |
| System ARS | Dr. Kumara | ISPV5_PRS_ID_1_7 | Dual Monitor Support and Span Mode | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_1_8 | Resizable Windows | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_1_9 | Multi-Resolution Support | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_2_2 | Support for lossy compression | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_2_6 | Copy Images to system (windows) clipboard | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_7_4 | Presentation State Support | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_7_5, ISPV5_PRS_ID_7_18 | Pre-Processing and Background Processing | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_7_12 | Launch Applications from another Application | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_7_25 | Web-based Training (on demand) | Dr. Kumara | TBD |
|  |  | ISPV5_PRS_ID_7_27 | Loading multiple studies into applications | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_9_1 to ISPV5_PRS_ID_1_3 | Languages and Localization | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_2_1, ISPV5_PRS_ID_7_10 | Save/Load Bookmarks support for clinical and viewing applications | Dr. Kumara | Done |
|  |  | ISPV5_PRS_ID_7_28 | Data Pre-fetch | Kishan Harwalkar | Done |
|  |  | ISPV5_PRS_ID_7_21 | PACS Integration | Kishan Harwalkar | Done |
|  |  | ISPV5_PRS_ID_7_22 | Desktop integration with iSITE PACS | Kishan Harwalkar | Done |
|  |  | ISPV5_PRS_ID_7_23 | Desktop integration with 3rd party PACS Vendors | Kishan Harwalkar | Done |
| Patient Directory ARS | Dr. Kumara | ISPV5_PRS_ID_2_7, ISPV5_PRS_ID_7_11 | Application common settings in preferences | ---- | ---- |
|  |  | ISPV5_PRS_ID_7_1, ISPV5_PRS_ID_7_2 | Patient Directory and Enhancements to PD |  |  |
|  |  | ISPV5_PRS_ID_7_6, ISPV5_PRS_ID_7_8 | CD/DVD Device Support and CD/DVD Burning |  |  |
|  |  | ISPV5_PRS_ID_7_7 | Provide a basic DICOM viewer application on burnt CD/DVD (media) |  |  |
|  |  | ISPV5_PRS_ID_7_9 | USB Device Support |  |  |
| Reporting ARS | Dr. Kumara | ISPV5_PRS_ID_7_3 | Reporting Application | ---- | ---- |
| Filming ARS | Dr. Kumara | ISPV5_PRS_ID_7_29 | Filming Application | ---- | ---- |
|  |  | ISPV5_PRS_ID_2_4 | UI Harmonization (usability improvements in Filming) |  |  |
| Collaboration ARS | Noa | ISPV5_PRS_ID_14_1 to ISPV5_PRS_ID_14_7 | Collaboration | ---- | ---- |
| Connectivity ARS | Kishan | ISPV5_PRS_ID_2_3 | Multi-vendor support | ---- | Medium |
|  |  | ISPV5_PRS_ID_7_19, ISPV5_PRS_ID_7_20 | Handling of CT/NM/MR Data and Enhanced DICOM Support |  |  |
| System Engineering ARS | Elan | ISPV5_PRS_ID_7_24 | High Availability | ---- | ---- |
|  |  | ISPV5_PRS_ID_7_26, ISPV5_PRS_ID_13_1 to ISPV5_PRS_ID_13_4 | Performance and Reliability | ---- | ---- |
|  |  | ISPV5_PRS_ID_1_12, ISPV5_PRS_ID_7_13, ISPV5_PRS_ID_7_14 | Licensing | ---- | ---- |
|  |  | ISPV5_PRS_ID_8_1 to ISPV5_PRS_ID_8_6 | Administration tools | ---- | ---- |
|  |  | ISPV5_PRS_ID_10_1 to ISPV5_PRS_ID_10_8 | Installation | ---- | ---- |
|  |  | ISPV5_PRS_ID_12_1 to ISPV5_PRS_ID_12_3 | Serviceability |  |  |
|  |  | ISPV5_PRS_ID_1_11 | Thin and Thick Client Support |  |  |
|  |  | ISPV5_PRS_ID_7_17 | Extended Storage |  |  |
| Security ARS | Sacchin Kumar N J | ISPV5_PRS_ID_10_1 to ISPV5_PRS_ID_10_6 | Security and Privacy | ---- | ---- |

## User Characteristics

TBD

## Assumptions and Dependencies

< These could include third-party or commercial components that you plan to use, issues around the clinical use and others. The project could be affected if these assumptions are incorrect, are not shared, or change. >

TBD

## References

### Internal Process Documents

< MRS…PRS

Any documents developed internally in the generation of this document, white papers, etc.>

### Standards Documents

<Any standards documents such as style guides, etc. >

# Requirements

## Dual Monitor Support and Span Mode (ISPV5_PRS_ID_1_7)

### This is not a new requirement for ISP 5.0. ISP 4.0 also supported dual monitor mode and SPREAD MODE. The following requirements are written to clarify certain additional aspects over the ISP 4.0 requirements. The ISP 4.0 System ARS should be used as a baseline for the fundamental requirements in this regard.

### System will support spanning over 2 screens. This option will be available whenever 2 or more screens are connected to the client (not be protected by license key). The Span mode is also known as SPREAD MODE and has been supported in ISP 4.0 also. The only addition to this requirement in ISP 5.0 is that MMV will support SPREAD MODE and NM Comprehensive Review application will support SPREAD MODE. No other MR Applications or NM applications including 3rd party NM Applications will support SPREAD MODE. All CT applications including review applications will continue to support SPREAD MODE as they did in ISP 4.0.

### The dual monitor support mode is not a new requirement and is the same as was defined for ISP 4.0. The System ARS for ISP 4.0 describes in detail about the dual monitor support requirements. All applications including CT, MR and Philips NM applications will support dual monitor configurations. However, NM 3rd party applications do not support dual monitor.

### For PIB (Portal in a Box), both monitors in dual monitor mode will have the same size, aspect ratio and resolution (including DPI).

### For Client-Server portal configuration, the monitors may be of different resolutions and the requirement is as detailed in the ISP 4.0 System ARS.

### The process to set up dual monitors on the system will remain the same as per the implementation in ISP 4.0

## Resizable Windows (ISPV5_PRS_ID_1_8)

### In ISP 5.0 client-server configuration, all viewer and application (including CT, MR and NM except for 3rd party NM applications) windows should support the ability to be minimized, maximized and also be able to be switched across screens.

## Multi-Resolution Support (ISPV5_PRS_ID_1_9)

### Multiple resolution support is not a new requirement for ISP 5.0. The ISP 4.0 System ARS defines the requirements as they were implemented in ISP 4.0 Additional relevant points are provided here.

### Portal Client shall support 2 resolutions (two ends of the spectrum):

- Low resolution 1024 X 768 and

- High Resolution 1280 X 1024

The minimum resolution supported is the Low Resolution and the maximum resolution supported is the High Resolution.

### The following resolutions should work without hitches:

- 1440x900

- 1366x768

- 1680x1050

- 1280X800

### Resolution should be determined when the application is launched. Run time adjustment to changes in resolution is not mandatory.

### The System UI will adjust itself to the screen resolution at the Client Side. The adjustments will be along the following lines:

- Workflow bar horizontal size shall adjust itself

- Workflow vertical size shall not change

- Toolbox horizontal size shall not change

- Toolbox vertical size shall adjust itself – buttons will be added and removed according to size. When removed, they will move into a ‘More’ section where they can be accessed

- Viewing area shall adjust both vertically and horizontally

- An example from CTV is shown below for reference:

(Note: Currently the DPI supported is ONLY 96 DPI. We need to consider making the system scale seamlessly to higher DPI settings (120, 200, 250 and 300), especially with the increasing availability of very high resolution and high DPI monitors which natively support a high DPI (example 300 DPI).

## Support for lossy compression (ISPV5_PRS_ID_2_2)

### This is not a new requirement – it is the same requirement as implemented in ISP 4.0. However the scope of support is defined here.

### The system shall allow and support the following compression levels:

- Lossless – no compression

- Lossy 80% - low or minimum compression

- Lossy 60% - medium or moderate compression

- Lossy 40% - high or maximum compression

### Compression will only affect image display.

### Saved images will have no compression and therefore ensure full quality.

### The system will clearly display to the user in terms of a message (status bar, header etc.) that the images being viewed are compressed images if any forms of compression is applied.

### Measurements on images will always be done using the lossless data

### Support for compression will be across all CT, MR and NM applications.

## Copy Images to Client’s Clipboard (ISPV5_PRS_ID_2_6)

### This is a requirement which was implemented in ISP 4.0 – but only for CT applications. The scope of this requirement is now expanded into MR Apps. (NOTE: NM Apps Scope is to be confirmed - TBD)

### All portal clients will allow the user to copy images to the system clipboard.

### The keyboard shortcut of [Ctrl+C] will be assigned to copy the images (selected images) to the windows clipboard.

### The image will be copied using a lossless rendering into the clipboard using a common format (bmp/png etc.) and user will be able to paste it into software that supports it. In ISP5.0, only CT and MR applications will have this feature. NM Support is to be confirmed.

## Presentation State Support (ISPV5_PRS_ID_7_4)

### Presentation states shall be supported in all viewing and clinical applications – CT and MR Applications. NM Support is yet to be confirmed.

### It should be possible to save presentation states.

### The system should be able to load saved presentation states.

### When images are sent to filming application, it should be possible to apply presentation states and send the images to filming; filming should support filming images that contain presentation states (applied presentation states).

### When images are sent to reporting application, it should be possible to apply presentation states and send the images to the reporting application; reporting should support adding images containing applied presentation states into summary images and then onto the report.

## Pre-processing and Background Processing (ISPV5_PRS_ID_7_5), (ISPV5_PRS_ID_7_18)

### Pre-Processing mechanism should be supported to enable automatic data processing based on configurable fields – such as Protocol Name, Body Part, Series Description and Procedure Description.

### In EX configuration, Pre-processing should be able to run on both the master and slave servers.

### Manual Configuration of Pre-Processing: This should be possible within preferences and the following parameters are supported:

- Algorithm Name

- Protocol Name

- Procedure Description

- Series Description

- Body Part

### When a field is empty, the system will not use the field in order to choose proper series

### It shall not be allowed to have all fields empty

### For all test fields the system will test whether the defined value is contained in the DICOM field.

### System will not look for exact match and will disregard upper/lower case

### Algorithm name shall display a pull-down of all available preprocessing algorithms

### Pre-processing will work if the “Pre-processing License” is valid.

### UI mockup of table

| Preprocessing | Procedure | Series Description | Protocol name | Body Part |
| --- | --- | --- | --- | --- |
| Liver segmentation ▼ | Abdomen | Portal | Liver | Liver |

### Automatic Configuration of Pre-Processing: When user that has permission to change system wide settings runs a manual preprocessing, a dialog will open in PD asking whether to add the series to the automatic preprocessing list. If the user approves, then the algorithm will be added with all the DICOM fields filled according to the series. {The Scope of this requirement is yet to be confirmed}

NOTE: The following are some important points to consider for future releases – these will be added to the appendix of this document:

### Support pre-processing algorithms that receive multiple series as an input either from the same study or from multiple studies (belonging to the same patient) and create one or more result series as the output (Liver, Tumor Tracking). Currently these applications do not support multiple series inputs. Hence this requirement should be considered for future versions of IntelliSpace Portal.

- Pre-processing tasks shall be notifiable – that is applications should be able to get a notification based on the processing tasks. Currently as of ISP 5.0, the architecture does not support accurate progress notifications. Only start and end points are notified. Hence we need to consider this requirement of having detailed and step wise accurate progress notification in future versions of IntelliSpace Portal.

- Automatic Result Creation: This requirement concerns the adoption of Closed Loop Imaging into the IntelliSpace Portal workflow and the enabling of CLI functions including the Pre-processing steps required in a CLI workflow. Since CLI integration is not in the scope of ISP 5.0, we need to consider this requirement for future versions of IntelliSpace Portal.

## Launch Applications from another application (ISPV5_PRS_ID_7_12)

### This is not an entirely new requirement. A part of this requirement with reference to CT Applications has already been implemented in ISP 4.0. The following points detail the requirements from NM perspective, which is newly added in ISP 5.0. The requirements are based on similar requirements detailed for EBW NM 2.0.

### It shall be possible to launch an NM application from a running NM application. When launching, all loaded data plus all data created in the running application shall be loaded into the new application. Launching of the new application shall be done from a button in the NM applications.

### The above requirement is for launching – Review Application from Analysis Application, Review Application from Review Application, Analysis Application from Analysis Application and Analysis Application from Review Application.

### Launching Application from Application (LAFA) will work for CT Applications in the same way as it is implemented in ISP 4.0; click on Review  opens CTV, or switching to currently running review application; click on Analysis  open application list (including NM applications), or switching to currently running analysis application.

### If launching application from application requires saving data to the disk, this data shall be saved into a temporary folder and not to a folder visible to the user. The data in the temporary folder shall be periodically deleted in order to free disk space.

## Web based training (On Demand) (ISPV5_PRS_ID_7_25) (TBD) – will be done shortly

## Loading multiple studies into an application (ISPV5_PRS_ID_7_27)

### The system shall allow multiple datasets (studies) to be loaded into an application.

### The applications will make use of this feature to optimize workflow within the application.

### In NM Applications, the following behavior is expected: The selections of data to be launched into the application supported are:

  - Single Study

  - Sub selection of Series

  - Sub selection of Images

  - Multiple Studies

  - Multiple patients

## Languages and Localization (ISPV5_PRS_ID_9_1 to ISPV5_PRS_ID_9_3)

### All clinical applications and viewers shall be localized (GUI) to the following language: (ISPV5_PRS_ID_9_1)

- English (United States)

- Danish (Denmark)

- Dutch (Netherlands)

- French (France)

- German (Germany)

- Italian (Italy)

- Norwegian (Norway)

- Spanish (Spain)

- Swedish (Sweden)

- Russian (Russian)

### The exceptions to the localization are:

- FAT Assessment – support only for English and Japanese.

- 3rd Party Nuclear Medicine (NM) applications (languages supported TBD).

### DICOM localization will be covered in detail in the Connectivity ARS. (ISPV5_PRS_ID_9_2)

### Localization of Report Templates ((ISPV5_PRS_ID_9_3)

      - The user shall be able to choose the following languages separately from the GUI language for report templates:

- Chinese

- Danish

- Dutch

- English

- French

- German

- Italian

- Japanese

- Norwegian

- Portuguese

- Spanish

- Swedish

- Russian

## Save/Load Bookmarks support for clinical and viewing applications (ISPV5_PRS_ID_2_1, ISPV5_PRS_ID_7_10)

### All CT applications and viewers in IntelliSpace Portal 5.0 will support saving/loading of application state using the concept of bookmarks.

### Bookmarks will not be supported by MR Applications (Including MMV) and NM Applications.

### Bookmarks are “pointers” to the state of an application. In any application or viewer (except the ones mentioned earlier), it should be possible to save the state of an application at any given point in time as a “bookmark”.

### The application should save the entire state in which the images/studies are at that point in time including any results created therein.

### It should be possible to load the application and restore it to the state it was when the bookmark was created by selecting the saved bookmark at any point of time later.

### Bookmarks should be accessible from the Patient Directory.

### A single bookmark can comprise of a number of “states” within it. When the bookmark is viewed in the patient directory window, it should be possible to expand the bookmark to see the states (individual states) contained within the bookmark. When the user selects a state, the user should be able to load that state. The application which was used to create the bookmark should be launched and the state opened.

### Bookmarks should have generic information which can be shared across applications (what exactly constitutes this information needs to be discussed and detailed).

### The user should be able to delete bookmarks either collectively as all states within a given bookmark or individual states within a given bookmark. The system will however, allow the user to delete one bookmark at a time.

## Data Prefetch (ISPV5_PRS_ID_7_28)

[Part Authored by Kishan Harwalkar]

#### Main Functions of Prefetch:

      - Prefetch shall automatically retrieve the previous studies of new exams that are planned or arrived to the Portal system

      - System shall query 1 or more predefined remote devices (typically only PACS) for the previous studies. In this case Prefetch will copy the previous studies into the local folder to which the new study arrived

      - The decision which studies to Prefetch will be based on predefined rules, that are configurable by user

      - The Prefetch will not retrieve the studies if it already exists in the local folder into which the new studies are retrieved

      - System will retrieve entire studies when deciding to Prefetch

      - The prefetched studies will have a flag in the Patient Directory indicating that the study has been prefetched

      - The pre-fetch flag appears only when it is prefetched directly into a particular local folder. If it is copied to another folder, the flag doesn’t appear.

      - Pre-fetch shall not be triggered for the studies which are already pre-fetched

#### Prefetch Operation: Prefetch operation will be initiated in two ways:

      - Using a RIS worklist. When RIS connection is established, query will be fired on the RIS to fetch the scheduled exams based on the configured query criteria and the results will decide the studies which need Prefetch. While Prefetching based on RIS Worklist, the previous studies shall be added into the local database of the Portal.

      - When a new study, with study date from the last 5 days, is added into Portal database, the system will look for previous exams.

      - The Prefetch will be triggered only when the new study is added to the Portal database via a DICOM Store from a remote device.

      - While Prefetching based on Study arrival, the previous studies shall be copied into the same folder into which the new study was added.

#### Prefetch Configuration:

      - A “Prefetch” page will be added to Preferences

      - Prefetch page will be editable only for users with permission to change system-wide-settings

      - User will be able to turn Prefetch on/off

      - Prefetch will be configured as a set of rules. User will have the following option for the rules

- Create one or more rule

- See all the rules configured

- Select any of the rules and Edit it

- Select a rule and delete it

- Enable/Disable a rule

#### Prefetch from PACS:

      - PACS device: It shall be possible to configure a PACS for doing the pre-fetch via the Lan Config

      - The query executed on the PACS for Prefetch shall contain the following parameters

- Patient ID

- Modality

- Study Date (configurable as last 6 months, last 1 year, last 2 years, last 5 years all)

- Maximum number of previous studies to retrieve – configurable

      - Identification of studies requiring Prefetch:

##### If RIS is connected to Portal, User will be able to configure whether to use the RIS worklist for evoking Prefetch. This can be done via the Lan-Config

##### User will be able to configure the interval for fetching the RIS worklist (even numbers from 2 hours to 24 hours)

##### User will be able to configure whether to initiate Prefetch for every new study that is created in Portal or only for those that appear in the RIS worklist or for both

##### User will identify the list of studies requiring Prefetch based on the following configurable parameters:

- Study Date – configurable (today, today + tomorrow)

- Modality – configurable

- Body Part Examined – configurable

- Study Description – configurable

- Series Description – configurable

#### Prefetch in Queue Manager:

##### Prefetch operations will be listed in the Transfer Page tab in the Queue Manager

##### Prefetch retrieval requests will be displayed in the same manner as regular copy operation

##### Prefetch retrieval operations will run in the clinical administrator account and a normal user will not be able to cancel/pause these retrievals.

##### Prefetch retrievals run in lower priority as compared to the user triggered transfer requests. The priority will be as follows:

##### Priority 0: Manual store/retrieve

##### Priority 1: Manual Prefetch

##### Priority 2: Auto Prefetch

#### Deletion of Prefetched Studies:

##### When auto-cleanup is triggered all the pre-fetched studies purged as the first priority. After which the remaining rules for clean-up shall apply.

#### Manual Prefetch:

##### Pre-Fetch via Context Menu:

##### Patient Directory shall provide a context menu option to manually Prefetch the previous exams from the PACS

##### When the user selects pre-fetch option from the Context menu, the PD shall query all the PACS that are configured for Pre-Fetch and brings up a window displaying all the previous studies for the patient in reverse chronological order.

##### The manual pre-fetch shall use the Patient ID & Patient Name for querying the previous studies.

##### The pre-fetch selection window shall display the following columns, if provided by PACS

#### •	Study Date

#### •	Modality

#### •	Study Description

#### •	Body part examined

#### •	Protocol

##### User shall be able to select one or more studies or select all studies and click on Pre-Fetch button to initiate the Pre-Fetch operation.

## PACS Integration (ISPV5_PRS_ID_7_21), Desktop integration with iSITE PACS (ISPV5_PRS_ID_7_22) and Desktop Integration with 3rd Party PACS Vendors (ISPV5_PRS_ID_7_23) [author – Kishan Harwalkar]

### Portal shall have API based integration with iSite PACS

### It shall be possible to launch portal applications on iSite by selecting one or more series/study.

### User shall be able to select the series to be launched via the series selector when not all series are available on iSite.

### It shall be possible to launch the bookmarks (created by Portal) stored on iSite server. The iSite Client shall directly launch the Portal application which created the bookmark.

### The bookmark shall be launched with the right portal application irrespective of whether the referenced series are stored in iSite or Portal

### Fast data storage from iSite to Portal.

      - The transfer of DICOM Image data from iSite to Portal shall be faster than the normal DICOM transfer.

      - The transfer of DICOM Image data from iSite to Portal shall be at the rate of at least 40 images/Second.

### It shall be possible for the user to install Portal Client during the iSite Client installation.

### iSite Authentication

      - With iSite integration it shall be possible to authenticate iSite users using iSite authentication system.

###### It shall be possible to import a group of roles from iSite and assign specific permissions on the Portal .

      - iSite users shall be able to login to Portal Client using iSite authentication system. (Stretch Goal TBD).

      - Portal shall provide a mechanism to monitor the health of the applications, which can be invoked by iSite Heartbeat.

### CDP (TBD)

### Portal shall have API based integration with CDP PACS.

### User shall be able to select the series to be launched via the series selector when not all series are available on CDP.

### 3rd Party PACS Integration:

### API based tight integration shall be provided for the following 3rd party PACS (TBD)

- McKesson

- Sectra

- Agfa

- GE (TBD)

- Fuji (TBD)

### Utilize 3 monitors: 1 – PD & report ; 2 – Review ; 3 – Analysis ; Film can be displayed in 2 or 3 (user settings).

### Automatic storage of results to PACS.

### When portal applications are launched by selecting Study/Series from PACS Worklist (via the PACS Plug-in), the results generated shall be automatically stored on the PACS.

### The results shall be stored on Local folder of Portal as well, if the selected data existed only in the local folder of portal.

### It shall be possible to allow filming and printing using Portal filming application with the PACS Integration. Printing shall be done on to the printer configured in the Portal and not on the PACS.

### Filming shall be disabled in the iSite Integration

## Portal in a Box (Equivalent to Single Box Workstation)

### Concept of Portal in a Box (PIB)

Portal in a box is the single box workstation equivalent in IntelliSpace Portal. The behavior for such a configuration is similar to but not exactly the same as for Extended Brilliance Workspace 4.5. However, PIB will support multiple modalities and multiple vendors (similar to the support offered by the client server configuration of IntelliSpace Portal).

### Intended clinical use of PIB:

PIB is intended for clinical usage scenarios where, a full blown client-server configuration such as that offered by IntelliSpace Portal is not required or is not part of the workflow in the clinical site. Typically, small to medium diagnostic labs where typically each modality is represented by a single scanner, are the sites where PIB will be used to the maximum. However, in the process of expansion of such single user diagnostics labs, PIB should be capable of supporting up to but not exceeding 2 client users logged into PIB (which will act as the limited Portal Server).

### Requirements where PIB is similar to client-server mode of ISP 5.0:

Unless specified in the following points of exception, PIB will support all functional features supported by ISP 5.0 client-server mode. The following points will define the differences that exist between ISP 5.0 Client-Server mode and PIB. This means that PIB will be multi-modality compliant and multi-vendor compliant as per the requirements defined for ISP 5.0.

### It will not be possible to Minimize and Maximize the Windows running in PIB:

###### Portal in a Box is a single box workstation. The clinical user, who is not an administrative user,

###### will not need to interact with the Windows Shell or Explorer. Hence PIB will not allow users to interact with the windows shell/explorer through any means.

###### All minimize and maximize buttons on the various application/viewer windows in IntelliSpace Portal will be disabled/not available in PIB.

###### The windows shortcut of Win+D which will minimize or maximize the active window will also be disabled in PIB. The reason for this is that the non-admin user will not need to interact with any application other than PIB workflow.

###### All other operating system shortcuts/mechanisms like Save As dialogs etc. which allow users to interact with the windows shell/explorer will also be disabled in PIB.

### Automatic logout from Windows Operating System session on exiting PIB workflow (for non-admin clinical users):

###### Whenever a non-admin clinical user exits the PIB workflow by any means, such a user will be automatically logged out of the windows session also.

###### When a non-admin clinical user exits the PIB workflow, the system logs the user out of the windows session and displays the login screen.

### Login into PIB:

###### PIB login screen will be a customized version of the windows login screen.

###### The following customizations will be applied to the login screen to make it appear as a Philips Branded login screen;

- Philips Brand logo will be displayed as specified by the Usability and standards requirements of Philips Healthcare.

- The login screen should prompt the user for User Name and Password

- The user should be able to choose the GUI localization language from within the login screen (similar to what it is in ISP 5.0 client-server scenario). The keyboard and Reporting languages selection will happen through the preferences page as in ISP 5.0 client-server scenario.

- The login screen colors will be as per Philips UI Harmonization Standards.

###### In PIB, even if the machine on which the PIB system is running is not connected to an active network, still a user should be able to login into the system and launch the workflow. This is true because, PIB does not require the client-server scenario unless a client is connecting into PIB. If the user is logging in on the same box as the PIB system a network connection is NOT MANDATORY.

### License Check on PIB Launch:

###### License check is mandatory before launching PIB workflow.

###### If a valid license is not available, the system will inform the user about the error giving a simple and appropriate message – example message – “A valid license which is required to run Philips IntelliSpace Portal is not available on your system. Please check with your system administrator regarding the same”.

###### In case a valid license is not found, after the user is shown the appropriate message, the application should exit and no icon is displayed in the Windows Notification Area.

### Data Download Scenarios in PIB:

###### All messages concerning data transfer/download on the network which are appropriate to a client-server scenario should not be displayed in PIB. An example of such a scenario and corresponding message is displayed below – such messages which are specific to client-server architecture should not be shown in PIB. Such messages should be removed from all possible scenarios such as filming, reporting, batch file saving etc.

### Help Window in PIB:

###### Product Info button shall reflect correct product name and version. The text shown on the image shall be localized.

###### Operational manual: Only the product-specific operation manual will be loaded.

###### Online Help webpage shall be opened inside the Help window (not in an external web browser).

### Display of Logged in User information on the Main Screen of PIB:

###### The message that is displayed at the left top of the workflow screen which shows the user, who has logged into the system, will remain similar to ISP 5.0 client-server scenario.

###### In the notification tray icon, change the product name to reflect installed version.

### Pre-configured users in PIB:

###### There will be only 1 system administrator account in PIB and such an account will be pre-configured (factory setting). No other user will be able to add any other admin accounts on the system.

###### All other user accounts as available in EBW 4.5 will be available in PIB, with similar use case scenarios.

### Changes in Preferences setting – available options for Admin and Non-Admin users in PIB:

###### Since there is no concept of client-server or multiple users in PIB, the concepts such as making header/footer in film as default and mandatory does not arise. Such settings will be disabled or not available in PIB.

(An exhaustive list of such settings will be provided)

### User Management in PIB:

[TBD] – This needs to be discussed with System Engineer (Elan) and all relevant requirements will be defined.

# Other Nonfunctional Requirements

## Users Rules

<List which individuals or roles can perform which functions..>

## Safety Requirements

<List here all the possible hazards related to this application. It may include: image orientations, CT Values, patient info etc.>

## Database

<List any database interfaces if needed.>

## Regulatory

<Describe any corporate or regulatory policies that will limit the options available to the developers..

Add a reference to corporate or Regulatory Policies

## Service s and Manufacturing

<List any special considerations for service or manufacturing.>

## Site Adaptations

<List any special modifications needed when installing the software.>

## Event Logging

<List any special considerations for logging of events.>

# PRS/PMR Requirements   to   Use Cases/Requirements   Tractability

<Link all relevant keys from PRS to this document >

# Revision History

| Revision | Date | Changed by: | Reason for change |
| --- | --- | --- | --- |
| 1.0 | 5th August 2011 | Dr. Kumara Sanjaya | New for IntelliSpace Portal 5.0 |

# Appendix A: Glossary

<Define all the terms, acronyms and abbreviations necessary to properly interpret the ARS..>

# Appendix B: To Be Determined List

<Collect a numbered list of the TBD (to be determined) references that remain in the ARS so they can be tracked to closure.>